## Readme
INSTRUCTIONS FOR CALIBRATION WORKBOOK
Purpose
This worksheet supports linear, quadratic, cubic, and quartic regression models:
| Linear: | y = b0 + b1x+ e
| Quadratic: | y = b0 + b1x + b2x2 + e
| Cubic: | y = b0 + b1x + b2x2 + b3x3 +e
| Quartic: | y = b0 + b1x + b2x2 + b3x3 +b4x4 + e
The inputs are:
|  | x | the concentration of the certified calibration standard
|  | y | the measured instrument response
|  | e | a random effect due to variation of influence quantities
It is assumed that the concentrations of the calibration standards (x) have negligible uncertainties.
The output quantities are estimated parameters (bs) and related uncertainties.
The workbook then helps the user perform the following functions:
- | determine which model (linear, quadratic, etc.) is best.
- | determine the replication of unknowns needed for uncertainty control.
- | determine whether zero and span responses are acceptable.
- | estimate the concentration and expanded uncertainty of candidate standards
analyzed on the same day as the initial calibration or a subsequent day
Warning
EPA discourages the use of cubic and quartic models and believes that better fits of the data
can be obtained by performing multipoint calibrations over more limited concentration ranges
and by using straight-line or quadratic models. Inclusion of cubic and quartic models in the
spreadsheet is for experimental use or for situations in which there is a theoretical basis for
the use of such higher-order models. Analysts should be aware that apparent higher-order
calibration curves may be caused by artifacts such as inaccurate reference standards or
leaks in a gas dilution system. They should not use higher-order regression models to fit
multipoint calibration data that have inadequate precision and that should be fitted to lower-
order regression models. If analysts suspect that precision is inadequate, they should make
replicate measurements at each different concentration. Additionally, a multipoint calibration
should not change regression model orders from one month to the next.
Organization
The workbook consists of several worksheets, which are displayed as tabs at the bottom of
the screen. The functions of these worksheets are described below:
ReadMe (this sheet) - describes the workbook, explaining how to use the worksheets
Measurement Data - allows for user input of calibration and other analytical data
and includes statistical calculations for polynomial regression
Curves 1 - displays the calibration data, the best-fit line, and expanded uncertainties.
Residuals 1 - displays the difference between the observed responses and those
estimated by the best-fit calibration line.
Curves 2 - displays the calibration data, the best-fit quadratic curve, and
expanded uncertainties.
Residuals 2 - displays the difference between the observed responses and those
estimated by the best-fit quadratic regression line.
Curves 3 - displays the calibration data, the best-fit cubic curve, and
expanded uncertainties.
Residuals 3 - displays the difference between the observed responses and those
estimated by the best fit cubic regression line.
Curves 4 - displays the calibration data, the best-fit quartic curve, and
expanded uncertainties.
Residuals 4 - displays the difference between the observed responses and those
estimated by the best-fit quartic regression line.
Chart Data - includes the data used to create the Curves and Residuals charts
Conventions
The Measurement Data worksheet includes instructions that guide the user through the
steps in its use. The worksheet is also color coded to simplify use. Shaded cells that are
bordered in blue lines are for input of data. These cells are unprotected, but all other cells of
the Measurement Data worksheet are protected. The only other unprotected cell in the
workbook is cell F4 of the Chart Data worksheet. That cell controls the approximate level of
confidence associated with the expanded uncertainties displayed in the "Curves" charts.
Derived values and statements are colored red. These cells contain formulas and are
protected to prevent alteration.
Spreadsheet background colors indicate the order of the polynomial supported by the
calculations in the area.
Light green is used for the linear model.
Yellow is used for the quadratic model.
Gray color is used for the cubic model.
Light blue is used for the quartic model.
Use
The Measurement Data worksheet guides the user through six steps.
STEP 1 |  |  | Enter Calibration Data
In this step, up to 50 calibration points may be entered. Each calibration point has two parts, the
certified concentration of the calibration gas standard and the instrument response when testing
the standard. These values are entered in two columns. The spreadsheet performs
computations in cloumns I through P (linear), Q through X (quadratic), Y through AZ (cubic), and BA
and above (quartic).
STEP 2 |  |  | Review the Parameter Estimates
In this step, the user reviews the estimates of the intercepts (b0), slopes (b1) and other
parameters (b2, b3, and b4) for the four models, examines their expanded uncertainties and the
overall uncertainty of the fit (u2). The result of an F test indicates which of the models is best. The
linear model is recommended unless the quadratic or higher order model significantly reduces
the residual error.
STEP 3 |  | Review the Charts
In this step, the user reviews the charts named Curves 1, Residuals 1, Curves 2, Residuals 2, etc.
These charts help the user understand why one model performs better than the other.
STEP 4 |  | Assess Magnitude of Uncertainty
In this step, the user enters the assumed concentration of a candidate standard and selects a
replication number, r. Based on the calibration results, the worksheet estimates the expanded
uncertainty associated with an approximate level of confidence of 95% that would result from
measuring such a standard r times. The user can use this as a guide for deciding whether to
proceed with analysis, producing additional calibration points, or taking some corrective action.
STEP 4a |  | Check Uncertainty Associated with Dilution
Following completion of the multipoint calibration, the accuracy of the assay apparatus must be
checked to verify that the error associated with the dilution is not excessive. This accuracy
check involves the measurement of an undiluted or diluted check standard.
Make 3 or more discrete measurements of the undiluted or diluted check standard. Enter the
response values in the step 4a section of the Measurement Data worksheet. The spreadsheet
calculates the relative difference between the mean analyzer response and the corresponding
response from the multipoint calibration regression equation and the undiluted or diluted check
standard concentration. If the relative difference is greater than 1.0%, the dilution error is
considered to be excessive. Assays may not be conducted until the relative difference for a
subsequent accuracy check is less than or equal to 1.0%.
STEP 5 |  | Assay Candidate Standard on Same Day
In this step, the user enters the responses to a candidate standard that is tested on the same day
as the calibration of STEP 1. The worksheet provides an estimate of the candidate's
concentration and its expanded uncertainty associated with an approximate level of confidence of 95%.
The worksheet also indicates whether the uncertainty in responses is larger than expected (unacceptable.)
STEP 6 |  | Assay Candidate Standard on
Different day from Initial Calibration
In this step, the user enters the responses to a candidate standard that is tested on a different day
from the calibration of STEP 1. The worksheet first assesses the zero and span responses. If
the zero and span responses are acceptable, the user proceeds to enter the results from testing
a candidate standard. The results include those for zero and nonzero reference standards. The
table below specifies the minimum number of different reference standard concentrations needed for
each model, in addition to the zero standard. Failure to meet these requirements will result in serious
calculation errors and invalid results.
| Model | Minimum number of different nonzero reference standards required
| Linear | 1
| Quadratic | 2
| Cubic | 3
| Quartic | 4
Whenever possible the concentrations of the additional reference standards should be located at
points which correspond to large differences between the polynomial calibration curve and the
corresponding straight line between the zero gas and the highest-concentration standard.
The spreadsheet determines whether the regression curve has changed since the initial
calibration. The data are corrected for any change and the estimated concentration of the
candidate standard is provided together with its expanded uncertainty associated with an
approximate level of confidence of 95%.
The spreadsheet also determines whether the relative combined standard uncertainty of the mean
response is acceptable (less than 1% of the mean response). This additional check is meant to
guard against hysteresis or other errors that are not corrected by the spreadsheet's adjustments.
Reference: Appendix A, EPA Traceability Protocol Assay and Certification of Gaseous
Calibration Standards, USEPA, 1997.

## Measurement Data
This spreadsheet is for the EPA protocol gases calculation.
| STEP 1 |  | Enter Calibration Data | [5 empty] | Calculations - Linear Model | [6 empty] | Calculations - Quadratic Model | [7 empty] | Calculations - Cubic Model | [6 empty] | Calculations - Cubic Model | [6 empty] | Calculations - Cubic Model | [6 empty] | Calculations - Cubic Model | [6 empty] | Calculations - Quartic Model | [6 empty] | Calculations - Quartic Model | [5 empty] | Calculations - Quartic Model | [6 empty] | Calculations - Quartic Model | [6 empty] | Calculations - Quartic Model | [6 empty] | Calculations - Quartic Model
Enter the calibration data in the shaded spaces below. The first column (i)
simply counts the calibration points that you enter. The second column (X) is
for the certified concentrations of the calibration gas standards. The third
column (Y) is for the instrument responses corresponding to the calibration
standards. The number of points cannot exceed 50.
| i | Xi | Yi |  | Color Code |  |  |  | X'X | [6 empty] | X2 |  | X'X | [5 empty] | X3 |  | X'X | [25 empty] | X4 |  | X'X
| 1 | 0 | 0.2194 | [5 empty] | 19 | 85.5 | [5 empty] | 0 |  | 19 | 85.5 | 527.25 |  |  |  | 0 |  | 19 | 85.5 | 527.25 | 3655.125 | [22 empty] | 0 |  | 19 | 85.5 | 527.25 | 3655.125 | 27021.5625
| 2 | 0.5 | 0.7141 |  | red = derived value (protected) |  |  |  | 85.5 | 527.25 | [5 empty] | 0.25 |  | 85.5 | 527.25 | 3655.125 |  |  |  | 0.125 |  | 85.5 | 527.25 | 3655.125 | 27021.5625 | [22 empty] | 0.0625 |  | 85.5 | 527.25 | 3655.125 | 27021.5625 | 208037.5312
| 3 | 1 | 1.2885 | [12 empty] | 1 |  | 527.25 | 3655.125 | 27021.5625 |  |  |  | 1 |  | 527.25 | 3655.125 | 27021.5625 | 208037.5312 | [22 empty] | 1 |  | 527.25 | 3655.125 | 27021.5625 | 208037.5312 | 1647030.1406
| 4 | 1.5 | 1.9132 |  | blue = entered value (unprotected) |  |  |  | Det(X'X) | [6 empty] | 2.25 | [7 empty] | 3.375 |  | 3655.125 | 27021.5625 | 208037.5312 | 1647030.1406 | [22 empty] | 5.0625 |  | 3655.125 | 27021.5625 | 208037.5312 | 1647030.1406 | 13307929.3828
| 5 | 2 | 2.591 | [5 empty] | 2707.5 | [6 empty] | 4 |  | Det(X'X) | [5 empty] | 8 |  | Det(X'X) | [25 empty] | 16 |  | 27021.5625 | 208037.5312 | 1647030.1406 | 13307929.3828 | 109207686.0352
| 6 | 2.5 | 3.2866 |  | black = fixed text (protected) | [10 empty] | 6.25 |  | 2295621.5625 | [5 empty] | 15.625 |  | 11011063605.6099 | [25 empty] | 39.0625 | 1 | 2 | 3 | 4 | 5
| 7 | 3 | 4.1078 | [5 empty] | X'Xinverse | [6 empty] | 9 | [7 empty] | 27 | [27 empty] | 81 | 2 | X'Xinverse
| 8 | 3.5 | 4.9446 | [5 empty] | 0.19474 | -0.031579 | [5 empty] | 12.25 |  | X'Xinverse | [5 empty] | 42.875 |  | X'Xinverse | [25 empty] | 150.0625 | 3 | 0.74537 | -0.94931 | 0.35677 | -0.051314 | 0.0024964
| 9 | 4 | 5.8145 | [5 empty] | -0.031579 | 0.0070175 | [5 empty] | 16 |  | 0.38647 | -0.16692 | 0.015038 |  |  |  | 64 |  | 0.58168 | -0.46867 | 0.10116 | -0.0063796 | [22 empty] | 256 | 4 | -0.94931 | 1.9803 | -0.89428 | 0.1418 | -0.00733
| 10 | 4.5 | 6.723 | [12 empty] | 20.25 |  | -0.16692 | 0.10255 | -0.010615 |  |  |  | 91.125 |  | -0.46867 | 0.56899 | -0.14374 | 0.0098613 | [22 empty] | 410.0625 | 5 | 0.35677 | -0.89428 | 0.43832 | -0.072982 | 0.0038982
| 11 | 5 | 7.7284 | [5 empty] | X'Y | [6 empty] | 25 |  | 0.015038 | -0.010615 | 0.0011794 |  |  |  | 125 |  | 0.10116 | -0.14374 | 0.039175 | -0.0028145 | [22 empty] | 625 |  | -0.051314 | 0.1418 | -0.072982 | 0.012543 | -0.00068528
| 12 | 5.5 | 8.7566 | [5 empty] | 142.0161 | [6 empty] | 30.25 | [7 empty] | 166.375 |  | -0.0063796 | 0.0098613 | -0.0028145 | 0.00020848 | [22 empty] | 915.0625 |  | 0.0024964 | -0.00733 | 0.0038982 | -0.00068528 | 3.8071e-05
| 13 | 6 | 9.8013 | [5 empty] | 909.8948 | [6 empty] | 36 |  | X'Y | [5 empty] | 216 |  | X'Y | [25 empty] | 1296
| 14 | 6.5 | 10.8818 | [12 empty] | 42.25 |  | 142.0161 | [5 empty] | 274.625 |  | 142.0161 | [25 empty] | 1785.0625 |  | X'Y |  | b
| 15 | 7 | 12.0982 | [5 empty] | b |  | b' |  |  |  |  | 49 |  | 909.8948 | [5 empty] | 343 |  | 909.8948 | [25 empty] | 2401 |  | 142.0161 |  | 0.22056
| 16 | 7.5 | 13.3122 | [5 empty] | -1.0778 |  | -1.0778 | 1.9005 |  |  |  | 56.25 |  | 6463.08 | [5 empty] | 421.875 |  | 6463.08 | [25 empty] | 3164.0625 |  | 909.8948 |  | 0.92846
| 17 | 8 | 14.584 | [5 empty] | 1.9005 | [6 empty] | 64 | [7 empty] | 512 |  | 48559.4466 | [25 empty] | 4096 |  | 6463.08 |  | 0.13904
| 18 | 8.5 | 15.9238 | [12 empty] | 72.25 |  | b |  | b' |  |  |  | 614.125 |  | b | [25 empty] | 5220.0625 |  | 48559.4466 |  | -0.0069299
| 19 | 9 | 17.3271 | [5 empty] | Y'Y | [6 empty] | 81 |  | 0.19642 |  | 0.19642 | 1.0011 | 0.099935 |  | 729 |  | 0.19517 | [25 empty] | 6561 |  | 378182.0241 |  | 0.00038726
[9 empty] | 1584.6791 | [8 empty] | 1.0011 | [7 empty] | 1.003
[18 empty] | 0.099935 | [7 empty] | 0.099384
[9 empty] | df | [16 empty] | 4.0861e-05
[9 empty] | 17 | [8 empty] | t(0.95,df) | [7 empty] | t(0.95,df) | [27 empty] | t(0.95,df)
[18 empty] | 2.1199 | [7 empty] | 2.1314 | [27 empty] | 2.1448
t(0.95,df)
[9 empty] | 2.1098 | [8 empty] | b'X'Y = SSmodel | [7 empty] | b'X'Y = SSmodel | [27 empty] | b'X'Y = SSmodel
[18 empty] | 1584.6714 | [7 empty] | 1584.6714 | [27 empty] | 1584.6753
b'X'Y = SSmodel
[9 empty] | 1576.2036 | [8 empty] | SSregression | [7 empty] | SSregression | [27 empty] | SSregression
[18 empty] | 523.1675 | [7 empty] | 523.1675 | [27 empty] | 523.1715
[9 empty] | SSregression | [18 empty] | b' | [27 empty] | b'
[9 empty] | 514.6998 | [8 empty] | SSresidual | [7 empty] | SSresidual |  | 0.19517 | 1.003 | 0.099384 | 4.0861e-05 | [22 empty] | SSresidual |  | 0.22056 | 0.92846 | 0.13904 | -0.0069299 | 0.00038726
[18 empty] | 0.007769 | [7 empty] | 0.007761 | [27 empty] | 0.0038217
SSresidual
[9 empty] | 8.4755 | [8 empty] | s2 | [7 empty] | s2 | [27 empty] | s2
[18 empty] | 0.00048556 | [7 empty] | 0.0005174 | [27 empty] | 0.00027298
s2
[9 empty] | 0.49856 | [8 empty] | Fcrit(0.05) | [7 empty] | Fcrit(0.05) | [27 empty] | Fcrit(0.05)
[18 empty] | 2.3167 | [7 empty] | 2.3849 | [27 empty] | 2.463
Var-Covar
[9 empty] | 0.097088 | -0.015744 | [7 empty] | F | [7 empty] | F | [27 empty] | F
[9 empty] | -0.015744 | 0.0034987 | [7 empty] | 1026.7642 | [7 empty] | 0.93847 | [27 empty] | 1.8954
[9 empty] | n | [8 empty] | Preferred Model | [7 empty] | Preferred Model | [27 empty] | Preferred Model
[9 empty] | 19 | [8 empty] | Q | [7 empty] | Quadratic | [27 empty] | Cubic
[18 empty] | Var-Covar | [7 empty] | Var-Covar | [27 empty] | Var-Covar
[18 empty] | 0.00018765 | -8.1049e-05 | 7.3017e-06 | [5 empty] | 0.00028244 | -0.00022757 | 4.9121e-05 | -3.0977e-06 | [24 empty] | 0.00020347 | -0.00025914 | 9.7392e-05 | -1.4008e-05 | 6.8146e-07
[18 empty] | -8.1049e-05 | 4.9795e-05 | -5.1542e-06 | [5 empty] | -0.00022757 | 0.00027628 | -6.9796e-05 | 4.7883e-06 | [24 empty] | -0.00025914 | 0.00054058 | -0.00024412 | 3.8709e-05 | -2.001e-06
[18 empty] | 7.3017e-06 | -5.1542e-06 | 5.7268e-07 | [5 empty] | 4.9121e-05 | -6.9796e-05 | 1.9022e-05 | -1.3666e-06 | [24 empty] | 9.7392e-05 | -0.00024412 | 0.00011965 | -1.9923e-05 | 1.0641e-06
[26 empty] | -3.0977e-06 | 4.7883e-06 | -1.3666e-06 | 1.0123e-07 | [24 empty] | -1.4008e-05 | 3.8709e-05 | -1.9923e-05 | 3.4241e-06 | -1.8707e-07
| STEP 2 |  | Review the Parameter Estimates | [50 empty] | 6.8146e-07 | -2.001e-06 | 1.0641e-06 | -1.8707e-07 | 1.0393e-08
Review the estimates of the parameters (b0, b1, …)
for the linear and quadratic models, their expanded
uncertainties and the overall uncertainty of the fit (u2).
The approximate level of confidence
| Linear Model |  | Expanded Uncertainty |  |  | associated with the expanded
| parameter | estimate | lower | upper |  | uncertainties is 95%.
| b0 = | -1.0778 | -1.7351 | -0.42036
| b1 = | 1.9005 | 1.7757 | 2.0253 |  | The corresponding intervals take the form
| s2 = | 0.49856 | 0.28073 | 1.1205 |  | estimate +/- t*u, where:
| s = | 0.70609 | 0.52984 | 1.0585 |  | estimate is the quantity of interest,
| df = | 17 |  |  |  | u is its uncertainty, and
| t = | 2.1098 |  |  |  | t is a coverage factor.
| Quadratic Model |  | Expanded Uncertainty |  |  | For each model,
| parameter | estimate | lower | upper |  | the coverage factor, t, is
| b0 = | 0.19642 | 0.19602 | 0.19682 |  | determined from the t-distribution
| b1 = | 1.0011 | 1.001 | 1.0012 |  | with appropriate degrees of
| b2 = | 0.099935 | 0.099934 | 0.099936 |  | freedom (df).
| s2 = | 0.00048556 | 0.00026799 | 0.0011247
| s = | 0.022036 | 0.01637 | 0.033537 |  | Note that the uncertainties are not
| df = | 16 |  |  |  | displayed, but can easily be
| t = | 2.1199 |  |  |  | derived from the estimate, expanded
uncertainty, and the coverage factor.
Comparing the two models:
| F ratio = | 1026.7642
| F critical = | 2.3167 | (5% significance level)
The quadratic model produces a significantly smaller overall uncertainty of fit.
The quadratic model appears to be the better choice.
If cubic or quartic models are supported by compelling
scientific theory or data, then review the following estimates
for those models. Otherwise, go to Step 3.
| Cubic Model |  | Expanded Uncertainty |  |  | See above notes on calculation
| parameter | estimate | lower | upper |  | of expanded uncertainties.
| b0 = | 0.19517 | 0.15935 | 0.23099
| b1 = | 1.003 | 0.9676 | 1.0385
| b2 = | 0.099384 | 0.090087 | 0.10868
| b3 = | 4.0861e-05 | -0.0006373 | 0.00071902
| s2 = | 0.0005174 | 0.00028089 | 0.0012394
| s = | 0.022746 | 0.01676 | 0.035204
| df = | 15
| t = | 2.1314
Comparing quadratic and cubic models:
| F ratio = | 0.93847
| F critical = | 2.3849 | (5% significance level)
The overall uncertainties of fit are not significantly different at the 5% level.
The quadratic model appears to be a better choice than cubic.
| Quartic Model |  | Expanded Uncertainty
| parameter | estimate | lower | upper
| b0 = | 0.22056 | 0.18997 | 0.25116
| b1 = | 0.92846 | 0.8786 | 0.97833
| b2 = | 0.13904 | 0.11558 | 0.1625
| b3 = | -0.0069299 | -0.010899 | -0.0029611
| b4 = | 0.00038726 | 0.00016861 | 0.00060591
| s2 = | 0.00027298 | 0.00014555 | 0.00067897
| s = | 0.016522 | 0.012064 | 0.026057
| df = | 14
| t = | 2.1448
Comparing quadratic and cubic models:
| F ratio = | 1.8954
| F critical = | 2.463 | (5% significance level)
The overall uncertainties of fit are not significantly different at the 5% level.
The cubic model appears to be a better choice than quartic.
| STEP 3 |  | Review the Charts
View the charts named Curves 1 and Residuals 1. Curves 1 shows bands
illustrating expanded uncertainties for the estimated regression. Compare these
bands with those of the quadratic regression, Curves 2.
(Note: You can change the associated approximate confidence level by
changing the "p-value" in cell F4 of the worksheet named Chart Data.)
Residuals 1 shows how the calibration points deviated from the calibration
line. Look for a simple pattern (such as a quadratic curve) in the chart. If
such a pattern appears, the quadratic model may be better. View Residual 2,
the deviations from the best- fit quadratic curve. If Residual 2 effectively
removes the simple pattern observed in Residual 1 and if the magnitude of
the deviations has been significantly reduced (as evidenced by a reduction
in the uncertainty u2), then the quadratic model is superior. An F-test can be
run to determine if the two overall uncertainties of fit are significantly different.
| F = | 1026.7642 |  | Prob. of greater F = | 4.5104e-21
The quadratic model produces a significantly smaller overall uncertainty of fit.
The quadratic model appears to be the better choice.
If cubic or quartic models are supported by compelling
scientific theory or data, then view their Curves and Residuals charts.
| STEP 4 |  | Assess Magnitude of Uncertainty | [5 empty] | STEP 4 |  | Assess Magnitude of Uncertainty |  |  |  |  | STEP 4 |  | Assess Magnitude of Uncertainty | [5 empty] | STEP 4 |  | Assess Magnitude of Uncertainty | [25 empty] | STEP 4 |  | Assess Magnitude of Uncertainty
| Enter the concentration at which you would like to evaluate the uncertainty of | [7 empty] | Calculations - Linear Model | [6 empty] | Calculations - Quadratic Model | [7 empty] | Calculations - Cubic Model | [27 empty] | Calculations - Quartic Model
| estimation and prediction. Also enter r, the number of assays to be performed | [50 empty] | x | xb | x'V | x'Vx
| Increasing r tends to reduce the prediction uncertainty, but with diminishing | [7 empty] | x | xb | x'V | x'Vx |  |  |  | x | xb | x'V | x'Vx |  |  |  |  | x | xb | x'V | x'Vx | [24 empty] | 1 | 9.8017 | 1.2209e-05 | 4.3825e-05
| effect. | [7 empty] | 1 | 10.3253 | 0.002624 | 0.034112 |  |  |  | 1 | 9.8006 | -3.5778e-05 | 4.9007e-05 |  |  |  |  | 1 | 9.7999 | 1.6263e-05 | 7.7579e-05 | [24 empty] | 6 |  | -3.6143e-05
[9 empty] | 6 |  | 0.005248 |  |  |  |  | 6 |  | 3.2171e-05 | [5 empty] | 6 |  | -4.8272e-05 | [25 empty] | 36 |  | 1.6006e-05
|  | Concentration = | 6 | [12 empty] | 36 |  | -3.0066e-06 | [5 empty] | 36 |  | 1.9953e-05 | [25 empty] | 216 |  | -1.7979e-06
[24 empty] | 216 |  | -1.7007e-06 | [25 empty] | 1296 |  | 4.6767e-08
|  | r = | 3
[9 empty] | 95% Conf. Limits for Response | [6 empty] | 95% Conf. Limits for Response
| Review the estimated mean response (estimate that only takes | [7 empty] | Estimate | Lower | Upper |  |  |  |  | Estimate | Lower | Upper | [5 empty] | 95% Conf. Limits for Response | [27 empty] | 95% Conf. Limits for Response
| into account the calibration uncertainty), and the expanded uncertainty. | [7 empty] | 10.3253 | 9.9356 | 10.715 |  |  |  |  | 9.8006 | 9.7858 | 9.8155 | [5 empty] | Estimate | Lower | Upper | [25 empty] | Estimate | Lower | Upper
| Review the predicted mean response and its expanded uncertainty. | [22 empty] | 9.7999 | 9.7812 | 9.8187 | [25 empty] | 9.8017 | 9.7875 | 9.8159
| To satisfy the EPA protocol requirements for prediction, the relative expanded | [7 empty] | 95% Conf. Limits for Response. | [6 empty] | 95% Conf. Limits for Response.
| uncertainty associated with an approximate 95% level of confidence | [7 empty] | Prediction | Lower | Upper |  |  |  |  | Prediction | Lower | Upper | Slope |  |  |  |  | 95% Conf. Limits for Response. | [27 empty] | 95% Conf. Limits for Response.
| should be less than 1% in absolute value. | [7 empty] | 10.3253 | 9.3811 | 11.2695 |  |  |  |  | 9.8006 | 9.7698 | 9.8314 | 1.6007 |  |  |  |  | Prediction | Lower | Upper | Slope (near Conc.) | [24 empty] | Prediction | Lower | Upper | Slope (near Conc.)
| Estimates below are based on the quadratic model. | [7 empty] | Conc. | Lower | Upper |  |  |  |  | Conc. | Lower | Upper | Check (Assumes Linearity in Neighborhood) |  |  |  |  | 9.7999 | 9.7662 | 9.8337 | 2.2 | [24 empty] | 9.8017 | 9.7768 | 9.8266 | 2.1831
| Tab-Right to view estimates based on the other models. | [7 empty] | 6 | 5.5032 | 6.4968 |  |  |  |  | 6 | 5.986 | 6.014 | 5.9808 | 6.0192 |  |  |  | Conc. | Lower | Upper | [25 empty] | Conc. | Lower | Upper
[10 empty] | -0.082806 | 0.082806 | [5 empty] | -0.0023332 | 0.0023303 | [5 empty] | 6 | 5.9847 | 6.0153 | [25 empty] | 6 | 5.9886 | 6.0114
[5 empty] | Expanded Uncertainty |  | See notes in Step 2 above | [17 empty] | -0.0025533 | 0.0025533 | [26 empty] | -0.0019013 | 0.0019013
|  |  |  | Estimate | Lower | Upper | for comments on calculation | [44 empty] | 95% Uncertainty in Prediction = |  | 0.1901%
|  |  | Instrument Response = | 9.8006 | 9.7858 | 9.8155 | of expanded uncertainties. |  | x | xb | x'V | x'Vx |  |  |  | x | xb | x'V | x'Vx
[9 empty] | 1 | 4.5 | 0.050881 | 0.03481 |  |  |  | 1 | 4.5 | 1.4818e-06 | 5.1041e-05 |  |  |  |  | x | xb | x'V | x'Vx | [24 empty] | x | xb | x'V | x'Vx
[5 empty] | Expanded Uncertainty |  |  |  | 2.9349 |  | -0.0054759 |  |  |  |  | 3.2466 |  | 2.6289e-05 | [5 empty] | 1 | 4.5 | -4.4637e-05 | 7.3488e-05 | [24 empty] | 1 | 4.5 | -1.5045e-05 | 4.3603e-05
|  |  |  | Prediction | Lower | Upper | [9 empty] | 10.5407 |  | -3.3955e-06 | [5 empty] | 3.2463 |  | 9.7589e-05 | [7 empty] | p | q | r | a | b | m | q | Solution 1 | Solution 2 | Solution 3 |  |  |  | Best Solution | Indicator | Unique? | A | B | 3.2428 |  | 2.5444e-05 | [9 empty] | quartic equation coefficients (r.h.s.=0) |  |  |  | resolvent cubic equation coefficients
|  |  | Instrument Response = | 9.8006 | 9.7698 | 9.8314 | [17 empty] | 10.5383 |  | -2.3747e-05 | [7 empty] | 2432.2567 | 24547.4475 | -105353.9847 | -1947410.117 | 1045838751.2099 | 1611.3804 | 1.0412 | 3.2463 | -2422.104 | -13.399 | 3.2463 | 999999 | 999999 | 3.2463 | -87553701267168 | #NUM! | #NUM! | #NUM! | 10.5159 |  | 2.3015e-06 | [9 empty] | a | b | c | d | pp | qq | rr | aa | bb | Indicator | mm | q | Conc1 | Conc2 | Conc3 | A | B | y y | R | D | E | Conc1 | Conc2 | Conc3 | Conc4 | In calibration range? |  |  |  | Best
|  |  | Concentration = | 6 | 5.986 | 6.014 | [17 empty] | 34.2103 |  | 1.5075e-06 | [25 empty] | 34.1011 |  | -1.9058e-06 | [9 empty] | -17.8945 | 359.024 | 2397.5035 | -11050.4909 | -359.024 | 1299.8641 | -18079082.8868 | -41666.2218 | -21351490.4085 | 1.1129e+14 | 235.701 | #NUM! | #NUM! | #NUM! | #NUM! | 276.8754 | 50.1624 | 446.7125 | 12.9515 | #NUM! | 10.4899 | #NUM! | #NUM! | 3.2428 | -7.2471 | 0 | 0 | 3.2428 | 0 | 3.2428
|  |  |  | Relative expanded uncertainty of prediction = | 0.2332% | [46 empty] | 110.5838 |  | 1.5306e-07
[25 empty] | Mean Response = | 4.5 | [26 empty] | Mean Response = | 4.5
| STEP 4a |  | Check Uncertainty Associated with Dilution | [5 empty] | STEP 4a |  | Check Uncertainty Associated with Dilution |  |  |  |  | STEP 4a |  | Check Uncertainty Associated with Dilution | [5 empty] | STEP 4a |  | Check Uncertainty Associated with Dilution | [25 empty] | STEP 4a |  | Check Uncertainty Associated with Dilution
This step checks the accuracy of the apparatus to verify that the error associated
with dilution is not excessive.
If no certified standards are diluted in performing the multipoint calibration, skip to Step 5.
|  |  | Check | [7 empty] | Check standard concentration = | 6 | [7 empty] | Check standard concentration = | 6 |  |  |  |  | Check standard concentration = | 6 | [26 empty] | Check standard concentration = | 6
|  |  | Standard | Response
|  |  | 6 | 9.8 | [6 empty] | Mean analyzer response = | 9.85 | [7 empty] | Mean analyzer response = | 9.85 |  |  |  |  | Mean analyzer response = | 9.85 | [26 empty] | Mean analyzer response = | 9.85
|  |  | 6 | 9.85
|  |  | 6 | 9.9 | [6 empty] | Predicted response = | 10.3253 | [7 empty] | Predicted response = | 9.8006 |  |  |  |  | Predicted response = | 9.7999 | [26 empty] | Predicted response = | 9.8017
[11 empty] | Relative difference = | -4.603% | [7 empty] | Relative difference = | 0.5037% |  |  |  |  | Relative difference = | 0.5108% | [26 empty] | Relative difference = | 0.4929%
[12 empty] | Relative difference is excessive | [8 empty] | Relative difference is ok | [5 empty] | Relative difference is ok | [27 empty] | Relative difference is ok
|  |  | mean = | 9.85
|  |  | Cal. Response = | 9.8006
|  |  | RD = | 0.5037%
Relative difference is ok
| STEP 5 |  | Assay Candidate Standard on Same Day | [5 empty] | STEP 5 |  | Assay Candidate Standard on Same Day |  |  |  |  | STEP 5 |  | Assay Candidate Standard on Same Day | [5 empty] | STEP 5 |  | Assay Candidate Standard on Same Day | [25 empty] | STEP 5 |  | Assay Candidate Standard on Same Day
| Proceed with the analysis of candidate standards if their relative expanded uncertainty, | [7 empty] | Calculations - Linear Model | [6 empty] | Calculations - Quadratic Model | [7 empty] | Calculations - Cubic Model | [27 empty] | Calculations - Quartic Model | [11 empty] | quartic equation coefficients (r.h.s.=0) |  |  |  | resolvent cubic equation coefficients
| as estimated above, is less than 1%. Enter the responses from the | [13 empty] | Metrology t | [6 empty] | Metrology t | [7 empty] | Metrology t |  |  |  | p | q | r | a | b | m | q | Solution 1 | Solution 2 | Solution 3 |  |  |  | Best Solution | Indicator | Unique? | A | B | [6 empty] | Metrology t | [5 empty] | a | b | c | d | pp | qq | rr | aa | bb | Indicator | mm | q | Conc1 | Conc2 | Conc3 | A | B | yy | R | D | E | Conc1 | Conc2 | Conc3 | Conc4 | In calibration range? |  |  |  | Best
| repeated analyses of an individual candidate standard in the spaces | [7 empty] | 2.9349 |  | 95% Conf. Limits for Response |  |  |  | 2 | 3.2466 |  | 95% Conf. Limits for Response |  |  |  | 2 |  | 3.2463 | 4.5 | 95% Conf. Limits for Response |  |  |  | 2 |  |  |  | 2432.2567 | 24547.4475 | -105353.9847 | -1947410.117 | 1045838751.2099 | 1611.3804 | 1.0412 | 3.2463 | -2422.104 | -13.399 | 3.2463 | 999999 | 999999 | 3.2463 | -87553701267168 | #NUM! | #NUM! | #NUM! | 3.2428 | 4.5 | 95% Conf. Limits for Response |  |  |  | 2 | [5 empty] | -17.8945 | 359.024 | 2397.5035 | -11050.4909 | -359.024 | 1299.8641 | -18079082.8868 | -41666.2218 | -21351490.4085 | 1.1129e+14 | 235.701 | #NUM! | #NUM! | #NUM! | #NUM! | 276.8754 | 50.1624 | 446.7125 | 12.9515 | #NUM! | 10.4899 | #NUM! | #NUM! | 3.2428 | -7.2471 | 0 | 0 | 3.2428 | 0 | 3.2428
| provided below. | [7 empty] | 2.9354 |  | Estimate | Lower | Upper |  |  | 3.2473 |  | Estimate | Lower | Upper |  |  |  | 3.2469 | 4.501 | Estimate | Lower | Upper | [5 empty] | 2432.2567 | 24547.4475 | -105378.4581 | -1947410.117 | 1045838726.7364 | 1611.3804 | 1.0412 | 3.2469 | -2422.104 | -13.3996 | 3.2469 | 999999 | 999999 | 3.2469 | -87566498904288 | #NUM! | #NUM! | #NUM! | 3.2434 | 4.501 | Estimate | Lower | Upper | [7 empty] | -17.8945 | 359.024 | 2397.5035 | -11053.0731 | -359.024 | 1310.1931 | -18081964.3554 | -41655.8929 | -21353135.7665 | 1.1131e+14 | 235.6718 | #NUM! | #NUM! | #NUM! | #NUM! | 276.883 | 50.1486 | 446.7063 | 12.9513 | #NUM! | 10.4909 | #NUM! | #NUM! | 3.2434 | -7.2474 | 0 | 0 | 3.2434 | 0 | 3.2434
[9 empty] | 2.9343 |  | 4.5 | 4.1268 | 4.8732 |  |  | 3.246 |  | 4.5 | 4.4857 | 4.5143 |  |  |  | 3.2457 | 4.499 | 4.5 | 4.4829 | 4.5171 | [5 empty] | 2432.2567 | 24547.4475 | -105329.5112 | -1947410.117 | 1045838775.6833 | 1611.3804 | 1.0412 | 3.2457 | -2422.104 | -13.3984 | 3.2457 | 999999 | 999999 | 3.2457 | -87540903629632 | #NUM! | #NUM! | #NUM! | 3.2422 | 4.499 | 4.5 | 4.4868 | 4.5132 | [7 empty] | -17.8945 | 359.024 | 2397.5035 | -11047.9087 | -359.024 | 1289.5352 | -18076201.4181 | -41676.5507 | -21349845.0504 | 1.1127e+14 | 235.7302 | #NUM! | #NUM! | #NUM! | #NUM! | 276.8678 | 50.1762 | 446.7187 | 12.9518 | #NUM! | 10.4889 | #NUM! | #NUM! | 3.2422 | -7.2467 | 0 | 0 | 3.2422 | 0 | 3.2422
| Note: This step applies only to candidate standards that are assayed on the | [32 empty] | 2432.2567 | 24547.4475 | 4776.5067 | -1947410.117 | 1045948881.7012 | 1611.3804 | 1.0437 | -0.19849 | -2422.1228 | -9.9354 | 999999 | 999999 | 999999 | 999999 | -29961301306944 | #NUM! | #NUM! | #NUM! | [12 empty] | -17.8945 | 359.024 | 2397.5035 | 569.5481 | -359.024 | -45180.2917 | -5112474.0129 | -88146.3777 | -13947379.2422 | 2.3266e+13 | 342.8243 | #NUM! | #NUM! | #NUM! | #NUM! | 227.646 | 129.0694 | 476.39 | 14.0506 | #NUM! | 4.6098 | #NUM! | #NUM! | -0.24679 | -4.8566 | 0 | 0 | 0 | 0 | 0
| same day as the calibration. | [9 empty] | 95% Conf. Limits for Response. | [6 empty] | 95% Conf. Limits for Response. | [7 empty] | 95% Conf. Limits for Response. | [7 empty] | 2432.2567 | 24547.4475 | 4776.5067 | -1947410.117 | 1045948881.7012 | 1611.3804 | 1.0437 | -0.19849 | -2422.1228 | -9.9354 | 999999 | 999999 | 999999 | 999999 | -29961301306944 | #NUM! | #NUM! | #NUM! |  |  | 95% Conf. Limits for Response. | [9 empty] | -17.8945 | 359.024 | 2397.5035 | 569.5481 | -359.024 | -45180.2917 | -5112474.0129 | -88146.3777 | -13947379.2422 | 2.3266e+13 | 342.8243 | #NUM! | #NUM! | #NUM! | #NUM! | 227.646 | 129.0694 | 476.39 | 14.0506 | #NUM! | 4.6098 | #NUM! | #NUM! | -0.24679 | -4.8566 | 0 | 0 | 0 | 0 | 0
[11 empty] | Prediction | Lower | Upper |  |  |  |  | Prediction | Lower | Upper | [5 empty] | Prediction | Lower | Upper | Slope (near Conc.) |  |  |  |  | 2432.2567 | 24547.4475 | 4776.5067 | -1947410.117 | 1045948881.7012 | 1611.3804 | 1.0437 | -0.19849 | -2422.1228 | -9.9354 | 999999 | 999999 | 999999 | 999999 | -29961301306944 | #NUM! | #NUM! | #NUM! |  |  | Prediction | Lower | Upper | Slope (near Conc.) | [6 empty] | -17.8945 | 359.024 | 2397.5035 | 569.5481 | -359.024 | -45180.2917 | -5112474.0129 | -88146.3777 | -13947379.2422 | 2.3266e+13 | 342.8243 | #NUM! | #NUM! | #NUM! | #NUM! | 227.646 | 129.0694 | 476.39 | 14.0506 | #NUM! | 4.6098 | #NUM! | #NUM! | -0.24679 | -4.8566 | 0 | 0 | 0 | 0 | 0
| Below, enter the instrument responses for up to 10 repeated analyses of | [9 empty] | 4.5 | 3.6033 | 5.3967 |  |  |  |  | 4.5 | 4.4708 | 4.5292 | [5 empty] | 4.5 | 4.4686 | 4.5314 | 1.6496 |  |  |  |  | 2432.2567 | 24547.4475 | 4776.5067 | -1947410.117 | 1045948881.7012 | 1611.3804 | 1.0437 | -0.19849 | -2422.1228 | -9.9354 | 999999 | 999999 | 999999 | 999999 | -29961301306944 | #NUM! | #NUM! | #NUM! |  |  | 4.5 | 4.4768 | 4.5232 | 1.6116 | [6 empty] | -17.8945 | 359.024 | 2397.5035 | 569.5481 | -359.024 | -45180.2917 | -5112474.0129 | -88146.3777 | -13947379.2422 | 2.3266e+13 | 342.8243 | #NUM! | #NUM! | #NUM! | #NUM! | 227.646 | 129.0694 | 476.39 | 14.0506 | #NUM! | 4.6098 | #NUM! | #NUM! | -0.24679 | -4.8566 | 0 | 0 | 0 | 0 | 0
| a single candidate standard. | [9 empty] | Conc. | Lower | Upper |  |  |  |  | Conc. | Lower | Upper | [5 empty] | Conc. | Lower | Upper | [5 empty] | 2432.2567 | 24547.4475 | 4776.5067 | -1947410.117 | 1045948881.7012 | 1611.3804 | 1.0437 | -0.19849 | -2422.1228 | -9.9354 | 999999 | 999999 | 999999 | 999999 | -29961301306944 | #NUM! | #NUM! | #NUM! |  |  | Conc. | Lower | Upper | [7 empty] | -17.8945 | 359.024 | 2397.5035 | 569.5481 | -359.024 | -45180.2917 | -5112474.0129 | -88146.3777 | -13947379.2422 | 2.3266e+13 | 342.8243 | #NUM! | #NUM! | #NUM! | #NUM! | 227.646 | 129.0694 | 476.39 | 14.0506 | #NUM! | 4.6098 | #NUM! | #NUM! | -0.24679 | -4.8566 | 0 | 0 | 0 | 0 | 0
[11 empty] | 2.9349 | 2.4631 | 3.4067 |  |  |  |  | 3.2466 | 3.2289 | 3.2644 | [5 empty] | 3.2463 | 3.2273 | 3.2653 | [5 empty] | 2432.2567 | 24547.4475 | 4776.5067 | -1947410.117 | 1045948881.7012 | 1611.3804 | 1.0437 | -0.19849 | -2422.1228 | -9.9354 | 999999 | 999999 | 999999 | 999999 | -29961301306944 | #NUM! | #NUM! | #NUM! |  |  | 3.2428 | 3.2284 | 3.2572 | [7 empty] | -17.8945 | 359.024 | 2397.5035 | 569.5481 | -359.024 | -45180.2917 | -5112474.0129 | -88146.3777 | -13947379.2422 | 2.3266e+13 | 342.8243 | #NUM! | #NUM! | #NUM! | #NUM! | 227.646 | 129.0694 | 476.39 | 14.0506 | #NUM! | 4.6098 | #NUM! | #NUM! | -0.24679 | -4.8566 | 0 | 0 | 0 | 0 | 0
| Analysis |  | Estimated | [8 empty] | -0.16076 | 0.16076 | [5 empty] | -0.0054533 | 0.0054533 | [6 empty] | -0.0058574 | 0.0058574 | [5 empty] | 2432.2567 | 24547.4475 | 4776.5067 | -1947410.117 | 1045948881.7012 | 1611.3804 | 1.0437 | -0.19849 | -2422.1228 | -9.9354 | 999999 | 999999 | 999999 | 999999 | -29961301306944 | #NUM! | #NUM! | #NUM! |  |  |  | -0.0044399 | 0.0044399 | [7 empty] | -17.8945 | 359.024 | 2397.5035 | 569.5481 | -359.024 | -45180.2917 | -5112474.0129 | -88146.3777 | -13947379.2422 | 2.3266e+13 | 342.8243 | #NUM! | #NUM! | #NUM! | #NUM! | 227.646 | 129.0694 | 476.39 | 14.0506 | #NUM! | 4.6098 | #NUM! | #NUM! | -0.24679 | -4.8566 | 0 | 0 | 0 | 0 | 0
| Number | Response | Conc. | [5 empty] | mean = | 2.9349 | [5 empty] | mean = | 3.2466 | [6 empty] | mean = | 3.2463 | [26 empty] | mean = | 3.2428
| 1 | 4.5 | 3.2466 |  | NOTE: For Cubic and |  |  |  | stdev = | 0.00052618 | [5 empty] | stdev = | 0.00060606 | [6 empty] | stdev = | 0.00060622 | [26 empty] | stdev = | 0.00060082
| 2 | 4.501 | 3.2473 |  | Quartic Model estimates, |  |  |  | df = | 2 | [5 empty] | df = | 2 | [6 empty] | df = | 2 | [26 empty] | df = | 2
| 3 | 4.499 | 3.246 |  | view the Calculations in |  |  |  | F = | 5.5532e-07 | [5 empty] | F = | 0.00075646 | [6 empty] | F = | 0.00071028 | [26 empty] | F = | 0.0013224
| 4 |  |  |  | the shaded regions to |  |  |  | F sig? = | FALSE | [5 empty] | F sig? = | FALSE | [6 empty] | F sig? = | FALSE | [26 empty] | F sig? = | FALSE
| 5 |  |  |  | the right. |  |  |  | Pr{>F} = | 1 | [5 empty] | Pr{>F} = | 0.99924 | [6 empty] | Pr{>F} = | 0.99929 | [26 empty] | Pr{>F} = | 0.99868
| 6 | [7 empty] | 95% U = | 16.08% | [5 empty] | 95% U = | 0.5453% | [6 empty] | 95% U = | 0.5857% | [26 empty] | 95% U = | 0.444%
| 7 | [7 empty] | 0.17319 | = Portion of Uncertainty2 Due to Calibration | [5 empty] | 0.23975 | = Portion of Uncertainty2 Due to Calibration | [6 empty] | 0.29879 | = Portion of Uncertainty2 Due to Calibration | [26 empty] | 32.4% | = Portion of Uncertainty2 Due to Calibration
8
| 9 | [22 empty] | STEP 6 |  | Assay Candidate Standard on | [25 empty] | STEP 6 |  | Assay Candidate Standard on
| 10 | [24 empty] | Different day from Initial Calibration | [27 empty] | Different day from Initial Calibration
|  | mean = | 3.2466
|  | standard uncertainty = | 0.00060606 | [21 empty] | Zero and Span Checks - Relative Difference Calculations for Cubic Model | [27 empty] | Zero and Span Checks - Relative Difference Calculations for Cubic Model
|  | df = | 2 | [23 empty] | Zero Gas Results |  | Span Gas Results | [25 empty] | Zero Gas Results |  | Span Gas Results
|  | F = | 0.00075646 | [21 empty] | Cal. Resp. = |  | 0.19517 |  | 17.3022 | [23 empty] | Cal. Resp. = |  | 0.22056 |  | 17.3276
|  | F sig? = | FALSE | (The uncertainty is acceptable.)
|  | Pr{>F} = | 0.99924 | [21 empty] | Relative Difference (RD)= |  | 1.128% |  | -7.526% | [23 empty] | Relative Difference (RD)= |  | 1.273% |  | -7.662%
|  | 95% Uncertainty = | 0.5453% | (This is the relative expanded uncertainty.) | [22 empty] | RD is okay. |  | RD is excessive. | [25 empty] | RD is okay. |  | RD is excessive.
|  | 23.97% | = Portion of Variance Due to Calibration
| Estimated Concentration of Candidate Standard | [7 empty] | STEP 6 |  | Assay Candidate Standard on |  |  |  |  | STEP 6 |  | Assay Candidate Standard on | [5 empty] | STEP 6 |  | Assay Candidate Standard on | [25 empty] | STEP 6 |  | Assay Candidate Standard on
| 3.2466 | [9 empty] | Different day from Initial Calibration | [6 empty] | Different day from Initial Calibration | [7 empty] | Different day from Initial Calibration | [27 empty] | Different day from Initial Calibration
[9 empty] | X'X | [6 empty] | X'X | [7 empty] | X'X | [27 empty] | X'X
| Rel. Expanded Uncertainty |  | Portion of Variance Due to Calibration Uncertainty | [5 empty] | 19 | 85.5 | 0 | 0 |  |  |  | 19 | 85.5 | 527.25 | 0 | 0 | 0 |  |  | 19 | 85.5 | 527.25 | 3655.125 | 0 | 0 | 0 | 0 | [20 empty] | 19 | 85.5 | 527.25 | 3655.125 | 27021.5625 | 0 | 0 | 0 | 0 | 0
| 0.5453% |  | 23.97% | [5 empty] | 85.5 | 527.25 | 0 | 0 |  |  |  | 85.5 | 527.25 | 3655.125 | 0 | 0 | 0 |  |  | 85.5 | 527.25 | 3655.125 | 27021.5625 | 0 | 0 | 0 | 0 | [20 empty] | 85.5 | 527.25 | 3655.125 | 27021.5625 | 208037.5312 | 0 | 0 | 0 | 0 | 0
[9 empty] | 0 | 0 | 9 | 40.5 |  |  |  | 527.25 | 3655.125 | 27021.5625 | 0 | 0 | 0 |  |  | 527.25 | 3655.125 | 27021.5625 | 208037.5312 | 0 | 0 | 0 | 0 | [20 empty] | 527.25 | 3655.125 | 27021.5625 | 208037.5312 | 1647030.1406 | 0 | 0 | 0 | 0 | 0
| STEP 6 |  | Assay Candidate Standard on | [5 empty] | 0 | 0 | 40.5 | 303.75 |  |  |  | 0 | 0 | 0 | 9 | 40.5 | 303.75 |  |  | 3655.125 | 27021.5625 | 208037.5312 | 1647030.1406 | 0 | 0 | 0 | 0 | [20 empty] | 3655.125 | 27021.5625 | 208037.5312 | 1647030.1406 | 13307929.3828 | 0 | 0 | 0 | 0 | 0
|  |  | Different Day from Initial Calibration | [12 empty] | 0 | 0 | 0 | 40.5 | 303.75 | 2460.375 |  |  | 0 | 0 | 0 | 0 | 9 | 40.5 | 303.75 | 2460.375 | [20 empty] | 27021.5625 | 208037.5312 | 1647030.1406 | 13307929.3828 | 109207686.0352 | 0 | 0 | 0 | 0 | 0
[9 empty] | Det(X'X) | UpLeft | DnRight |  |  |  |  | 0 | 0 | 0 | 303.75 | 2460.375 | 20913.1875 |  |  | 0 | 0 | 0 | 0 | 40.5 | 303.75 | 2460.375 | 20913.1875 | [20 empty] | 0 | 0 | 0 | 0 | 0 | 9 | 40.5 | 303.75 | 2460.375 | 20913.1875
| This step applies to candidate standards that are assayed on a different day | [7 empty] | 2960651.25 | 2707.5 | 1093.5 | [12 empty] | 0 | 0 | 0 | 0 | 303.75 | 2460.375 | 20913.1875 | 182682.8438 | [20 empty] | 0 | 0 | 0 | 0 | 0 | 40.5 | 303.75 | 2460.375 | 20913.1875 | 182682.8438
| than the initial calibration. Before candidate standards are run, the | [14 empty] | Det(X'X) | UpLeft | DnRight | [5 empty] | 0 | 0 | 0 | 0 | 2460.375 | 20913.1875 | 182682.8438 | 1619234.2969 | [20 empty] | 0 | 0 | 0 | 0 | 0 | 303.75 | 2460.375 | 20913.1875 | 182682.8438 | 1619234.2969
| measurement system is challenged with zero and span checks. Three or | [7 empty] | X'Xinverse | [6 empty] | 2.0587e+12 | 2295621.5625 | 896806.6875 | [33 empty] | 0 | 0 | 0 | 0 | 0 | 2460.375 | 20913.1875 | 182682.8438 | 1619234.2969 | 14461007.8359
| more discrete checks of the zero gas and three or more checks of the span | [7 empty] | 0.19474 | -0.031579 | 0 | 0 | [11 empty] | X'Xinverse | [27 empty] | 0 | 0 | 0 | 0 | 0 | 20913.1875 | 182682.8438 | 1619234.2969 | 14461007.8359 | 129644616.7617
| gas are made. Enter the results below: | [7 empty] | -0.031579 | 0.0070175 | 0 | 0 |  |  |  | X'Xinverse | [7 empty] | 0.58168 | -0.46867 | 0.10116 | -0.0063796 | 0 | 0 | 0 | 0
[9 empty] | 0 | 0 | 0.27778 | -0.037037 |  |  |  | 0.38647 | -0.16692 | 0.015038 | 0 | 0 | 0 |  |  | -0.46867 | 0.56899 | -0.14374 | 0.0098613 | 0 | 0 | 0 | 0 | [20 empty] | X'Xinverse |  |  |  |  | 1 | 2 | 3 | 4 | 5
|  | Response to |  | Span | Response |  |  |  | 0 | 0 | -0.037037 | 0.0082305 |  |  |  | -0.16692 | 0.10255 | -0.010615 | 0 | 0 | 0 |  |  | 0.10116 | -0.14374 | 0.039175 | -0.0028145 | 0 | 0 | 0 | 0 | [20 empty] | 0.74537 | -0.94931 | 0.35677 | -0.051314 | 0.0024964 | 0 | 0 | 0 | 0 | 0
|  | Zero gas |  | Conc. | to Span | [10 empty] | 0.015038 | -0.010615 | 0.0011794 | 0 | 0 | 0 |  |  | -0.0063796 | 0.0098613 | -0.0028145 | 0.00020848 | 0 | 0 | 0 | 0 | [20 empty] | -0.94931 | 1.9803 | -0.89428 | 0.1418 | -0.00733 | 0 | 0 | 0 | 0 | 0
|  | 0 |  | 9 | 16.01 |  |  |  | X'Y | b | Y'Y | df |  | Estimate |  | 0 | 0 | 0 | 0.33333 | -0.11111 | 0.0082305 |  | Estimate | 0 | 0 | 0 | 0 | 0.33333 | 0 | -0.83239 | 0.052025 | [20 empty] | 0.35677 | -0.89428 | 0.43832 | -0.072982 | 0.0038982 | 0 | 0 | 0 | 0 | 0 |  |  | quartic equation coefficients (r.h.s.=0) |  |  |  | resolvent cubic equation coefficients
|  | 0.001 |  | 9 | 16 |  |  |  | 142.0161 | -1.0778 | 2621.3774 | 24 |  | 2.3525 |  | 0 | 0 | 0 | -0.11111 | 0.107 | -0.010974 |  | 2.94 | 0 | 0 | 0 | 0 | 1.8889 | -712483534798848 | 2.3749e+14 | -1.7592e+13 |  | Estimate | p | q | r | a | b | m | q | Solution 1 | Solution 2 | Solution 3 |  |  |  | Best Solution | Indicator | Unique? | A | B | -0.051314 | 0.1418 | -0.072982 | 0.012543 | -0.00068528 | 0 | 0 | 0 | 0 | 0 |  | Estimate | a | b | c | d | pp | qq | rr | aa | bb | Indicator | mm | q | Conc1 | Conc2 | Conc3 | A | B | yy | R | D | E | Conc1 | Conc2 | Conc3 | Conc4 | In calibration range? |  |  |  | Best
|  | -0.001 |  | 9 | 15.99 |  |  |  | 909.8948 | 1.9005 |  |  |  | 2.3473 |  | 0 | 0 | 0 | 0.0082305 | -0.010974 | 0.0012193 |  | 2.9337 | 0 | 0 | 0 | 0 | -0.65844 | 237494511599616 | -7.9165e+13 | 5.8641e+12 |  | 0.029837 | 16 | -64 | 1.8953 | -149.3333 | 646.636 | 14.1107 | 0.91348 | 3.2882 | -19.318 | 0.029837 | 3.2882 | 999999 | 0.029837 | 0.029837 | -18806.1875 | #NUM! | #NUM! | #NUM! | 0.0024964 | -0.00733 | 0.0038982 | -0.00068528 | 3.8071e-05 | 0 | 0 | 0 | 0 | 0 |  | 2.9432 | -17.8945 | 359.024 | 2397.5035 | -9785.1978 | -359.024 | -3761.3084 | -16667163.2538 | -46727.3943 | -20545264.9704 | 1.0175e+14 | 249.6061 | #NUM! | #NUM! | #NUM! | #NUM! | 273.0592 | 57.0418 | 449.7757 | 13.0692 | #NUM! | 10.0084 | #NUM! | #NUM! | 2.9432 | -7.0652 | 0 | 0 | 2.9432 | 0 | 2.9432
[9 empty] | 72.8285 | -0.46179 | Metrology t | b'X'Y |  | 2.342 | [8 empty] | 2.9274 | 0 | 0 | 0 | 0 | 0.049383 | -17592186044416 | 5.8641e+12 | -434374964059.6536 |  | 0.029758 | 16 | -64 | 1.8903 | -149.3333 | 646.631 | 14.1107 | 0.91347 | 3.2883 | -19.318 | 0.029758 | 3.2883 | 999999 | 0.029758 | 0.029758 | -18807.8041 | #NUM! | #NUM! | #NUM! | 0 | 0 | 0 | 0 | 0 | 0.33333 | 0 | -0.14689 | 0 | 0 | 1 | 2.937 | -17.8945 | 359.024 | 2397.5035 | -9759.3755 | -359.024 | -3864.5976 | -16638348.5675 | -46830.6836 | -20528811.39 | 1.0155e+14 | 249.8818 | #NUM! | #NUM! | #NUM! | #NUM! | 272.9794 | 57.1846 | 449.8387 | 13.0716 | #NUM! | 9.9984 | #NUM! | #NUM! | 2.937 | -7.0614 | 0 | 0 | 2.937 | 0 | 2.937
[9 empty] | 558.6822 | 1.9009 | 2 | 2604.546 |  |  |  | X'Y | b | Y'Y | df | [13 empty] | 0.029678 | 16 | -64 | 1.8853 | -149.3333 | 646.626 | 14.1107 | 0.91347 | 3.2883 | -19.318 | 0.029678 | 3.2883 | 999999 | 0.029678 | 0.029678 | -18809.4206 | #NUM! | #NUM! | #NUM! | 0 | 0 | 0 | 0 | 0 | -0.26963 | 1.6942e+14 | -2.8412e+13 | -5.1705e+12 | 692865429935.3926 | 2 | 2.9307 | -17.8945 | 359.024 | 2397.5035 | -9733.5532 | -359.024 | -3967.8869 | -16609533.8811 | -46933.9728 | -20512357.8096 | 1.0136e+14 | 250.1572 | #NUM! | #NUM! | #NUM! | #NUM! | 272.8995 | 57.3275 | 449.9017 | 13.0741 | #NUM! | 9.9883 | #NUM! | #NUM! | 2.9307 | -7.0576 | 0 | 0 | 2.9307 | 0 | 2.9307
[16 empty] | 142.0161 | 0.19642 | 2621.3774 | 22 |  |  |  |  | X'Y | b | Y'Y | df | [6 empty] | 16 | -64 | -0.1097 | -149.3333 | 644.631 | 14.1107 | 0.91106 | 3.3152 | -19.3135 | -0.0017133 | 3.3152 | 999999 | 999999 | 3.3152 | -19453.4351 | #NUM! | #NUM! | #NUM! | 0 | 0 | 0 | 0 | 0 | 0.048042 | -4.2548e+13 | 7.0591e+12 | 1.324e+12 | -175890025145.4913 | 3 |  | -17.8945 | 359.024 | 2397.5035 | 569.5481 | -359.024 | -45180.2917 | -5112474.0129 | -88146.3777 | -13947379.2422 | 2.3266e+13 | 342.8243 | #NUM! | #NUM! | #NUM! | #NUM! | 227.646 | 129.0694 | 476.39 | 14.0506 | #NUM! | 4.6098 | #NUM! | #NUM! | -0.24679 | -4.8566 | 0 | 0 | 0 | 0 | 0
[9 empty] | SSregr | SSresid | s2 | s |  |  |  | 909.8948 | 1.0011 | [6 empty] | 142.0161 | 0.19517 | 2621.3774 | 20 | [6 empty] | 16 | -64 | -0.1097 | -149.3333 | 644.631 | 14.1107 | 0.91106 | 3.3152 | -19.3135 | -0.0017133 | 3.3152 | 999999 | 999999 | 3.3152 | -19453.4351 | #NUM! | #NUM! | #NUM! | 0 | 0 | 0 | 0 | 0 | 0.00042895 | -458428300831.6879 | 102295768300.8209 | 5518671132.0944 | -1247250995.3023 | 4 |  | -17.8945 | 359.024 | 2397.5035 | 569.5481 | -359.024 | -45180.2917 | -5112474.0129 | -88146.3777 | -13947379.2422 | 2.3266e+13 | 342.8243 | #NUM! | #NUM! | #NUM! | #NUM! | 227.646 | 129.0694 | 476.39 | 14.0506 | #NUM! | 4.6098 | #NUM! | #NUM! | -0.24679 | -4.8566 | 0 | 0 | 0 | 0 | 0
[9 empty] | 956.0387 | 16.8315 | 0.70131 | 0.83744 |  |  |  | 6463.08 | 0.099935 | [6 empty] | 909.8948 | 1.003 | [8 empty] | 16 | -64 | -0.1097 | -149.3333 | 644.631 | 14.1107 | 0.91106 | 3.3152 | -19.3135 | -0.0017133 | 3.3152 | 999999 | 999999 | 3.3152 | -19453.4351 | #NUM! | #NUM! | #NUM! | 0 | 0 | 0 | 0 | 0 | -0.00032171 | 343821225623.7659 | -59541957041.6157 | -9865626410.4042 | 1359632547.3162 | 5 |  | -17.8945 | 359.024 | 2397.5035 | 569.5481 | -359.024 | -45180.2917 | -5112474.0129 | -88146.3777 | -13947379.2422 | 2.3266e+13 | 342.8243 | #NUM! | #NUM! | #NUM! | #NUM! | 227.646 | 129.0694 | 476.39 | 14.0506 | #NUM! | 4.6098 | #NUM! | #NUM! | -0.24679 | -4.8566 | 0 | 0 | 0 | 0 | 0
[16 empty] | 72.8285 | 0.2194 | Metrology t | b'X'Y |  |  |  |  | 6463.08 | 0.099384 | [8 empty] | 16 | -64 | -0.1097 | -149.3333 | 644.631 | 14.1107 | 0.91106 | 3.3152 | -19.3135 | -0.0017133 | 3.3152 | 999999 | 999999 | 3.3152 | -19453.4351 | #NUM! | #NUM! | #NUM! | [12 empty] | -17.8945 | 359.024 | 2397.5035 | 569.5481 | -359.024 | -45180.2917 | -5112474.0129 | -88146.3777 | -13947379.2422 | 2.3266e+13 | 342.8243 | #NUM! | #NUM! | #NUM! | #NUM! | 227.646 | 129.0694 | 476.39 | 14.0506 | #NUM! | 4.6098 | #NUM! | #NUM! | -0.24679 | -4.8566 | 0 | 0 | 0 | 0 | 0
[16 empty] | 558.6822 | 0.9926 | 2 | 2621.3662 |  |  |  |  | 48559.4466 | 4.0861e-05 | Metrology t | b'X'Y |  | x | x'V |  |  |  | 16 | -64 | -0.1097 | -149.3333 | 644.631 | 14.1107 | 0.91106 | 3.3152 | -19.3135 | -0.0017133 | 3.3152 | 999999 | 999999 | 3.3152 | -19453.4351 | #NUM! | #NUM! | #NUM! | X'Y | b | Y'Y | df | [8 empty] | -17.8945 | 359.024 | 2397.5035 | 569.5481 | -359.024 | -45180.2917 | -5112474.0129 | -88146.3777 | -13947379.2422 | 2.3266e+13 | 342.8243 | #NUM! | #NUM! | #NUM! | #NUM! | 227.646 | 129.0694 | 476.39 | 14.0506 | #NUM! | 4.6098 | #NUM! | #NUM! | -0.24679 | -4.8566 | 0 | 0 | 0 | 0 | 0
| n = | 3 |  | n = | 3 | [10 empty] | 4619.3126 | 0.10092 | [6 empty] | 72.8285 | 0.2194 | 2 | -154174.2107 |  | 1 | 3049.3655 |  |  |  | 16 | -64 | -0.1097 | -149.3333 | 644.631 | 14.1107 | 0.91106 | 3.3152 | -19.3135 | -0.0017133 | 3.3152 | 999999 | 999999 | 3.3152 | -19453.4351 | #NUM! | #NUM! | #NUM! | 142.0161 | 0.22056 | 2621.3774 | 18 | [8 empty] | -17.8945 | 359.024 | 2397.5035 | 569.5481 | -359.024 | -45180.2917 | -5112474.0129 | -88146.3777 | -13947379.2422 | 2.3266e+13 | 342.8243 | #NUM! | #NUM! | #NUM! | #NUM! | 227.646 | 129.0694 | 476.39 | 14.0506 | #NUM! | 4.6098 | #NUM! | #NUM! | -0.24679 | -4.8566 | 0 | 0 | 0 | 0 | 0
| mean = | 0 |  | mean = | 16 | [7 empty] | nnn = | 3 | [7 empty] | nnn = | 3 | 558.6822 | 128 |  |  |  | 0.029758 | -1.6457e+17 |  |  |  | 16 | -64 | -0.1097 | -149.3333 | 644.631 | 14.1107 | 0.91106 | 3.3152 | -19.3135 | -0.0017133 | 3.3152 | 999999 | 999999 | 3.3152 | -19453.4351 | #NUM! | #NUM! | #NUM! | 909.8948 | 0.92846 | [10 empty] | -17.8945 | 359.024 | 2397.5035 | 569.5481 | -359.024 | -45180.2917 | -5112474.0129 | -88146.3777 | -13947379.2422 | 2.3266e+13 | 342.8243 | #NUM! | #NUM! | #NUM! | #NUM! | 227.646 | 129.0694 | 476.39 | 14.0506 | #NUM! | 4.6098 | #NUM! | #NUM! | -0.24679 | -4.8566 | 0 | 0 | 0 | 0 | 0
| s = | 0.001 |  | s = | 0.01 | s = experimental standard deviation | [6 empty] | mean = | 2.3473 |  | SSregr | SSresid | s2 | s |  |  | mean = | 2.9337 | 4619.3126 | -32 |  |  |  | 0.00088552 | 54857678867093232 |  | nnn = | 3 | [18 empty] | 6463.08 | 0.13904 | [8 empty] | nnn = | 3
| Cal. Resp. = | 0.19642 |  | Cal. Response = | 17.301 | [7 empty] | stdev = | 0.0052608 |  | 972.859 | 0.011238 | 0.0005108 | 0.022601 |  |  | stdev = | 0.0063102 | 39734.0903 | -2 |  |  |  | 2.6351e-05 | -4063531767932905 |  | mean = | 0.029758 | 16 | -64 | 1.8903 | -149.3333 | 646.631 | 14.1107 | 0.91347 | 3.2883 | -19.318 | 0.029758 | 3.2883 | 999999 | 0.029758 | 0.029758 | -18807.8041 | #NUM! | #NUM! | #NUM! | 48559.4466 | -0.0069299 | Metrology t | b'X'Y |  | x | x'V |  |  |  | mean = | 2.937 | -17.8945 | 359.024 | 2397.5035 | -9759.3755 | -359.024 | -3864.5976 | -16638348.5675 | -46830.6836 | -20528811.39 | 1.0155e+14 | 249.8818 | #NUM! | #NUM! | #NUM! | #NUM! | 272.9794 | 57.1846 | 449.8387 | 13.0716 | #NUM! | 9.9984 | #NUM! | #NUM! | 2.937 | -7.0614 | 0 | 0 | 2.937 | 0 | 2.937
[13 empty] | df = | 2 | [7 empty] | df = | 2 | [8 empty] | stdev = | 7.9308e-05 | [18 empty] | 378182.0241 | 0.00038726 | 2 | -5971.7225 |  | 1 | -27.3272 |  |  |  | stdev = | 0.0062303
|  |  | Zero Gas Results |  | Span Gas Results | [7 empty] | F = | 3.9463e-05 | [7 empty] | F = | 0.077954 | SSregr | SSresid | s2 | s |  |  |  |  | df = | 2 | [18 empty] | 72.8285 | 0.2194 |  |  |  | 2.937 | 68999567557077816 |  |  |  | df = | 2
|  | Std. Uncertainty = s/sqrt(n) = | 0.00057735 |  | 0.0057735 | [7 empty] | F sig? = | FALSE | [7 empty] | F sig? = | FALSE | -155822.7179 | 156795.5881 | 7839.7794 | 88.5425 |  |  |  |  | F = | 8.0229e-13 | [18 empty] | 558.6822 | 16 |  |  |  | 8.6259 | -11644781843045350 |  |  |  | F = | 8.1308e-08
|  | Rrs/100 = | 0.16 |  | 0.16 | [7 empty] | Pr{>F} = | 0.99996 | [7 empty] | Pr{>F} = | 0.92526 | [8 empty] | F sig? = | FALSE | [18 empty] | 4619.3126 | -4 |  |  |  | 25.334 | -2.0813e+15 |  |  |  | F sig? = | FALSE
|  |  | Std. Uncertainty is okay. |  | Std. Uncertainy is okay. | [26 empty] | Pr{>F} = | 1 | [18 empty] | 39734.0903 | -0.5 |  |  |  | 74.4053 | 2.8037e+14 |  |  |  | Pr{>F} = | 1
|  | Relative Difference (RD)= | 1.135% |  | -7.52% | [26 empty] | std error = | 0.0015387 | [18 empty] | 349328.0611 | 0.0625 | [8 empty] | std error = | 0.0012247
|  |  | RD is okay. |  | RD is excessive. | [26 empty] | The relative standard uncertainty is okay. | [29 empty] | The relative standard uncertainty is okay.
[52 empty] | SSregr | SSresid | s2 | s
| RD calculations above are based on linear/quadratic model. Tab right to see results for cubic/quartic. | [50 empty] | -7620.2297 | 8593.0999 | 477.3944 | 21.8494
Following successful completion of the zero and span checks, the candidate
standard is measured together with reference standards. While the candidate
standard is normally interspersed with the reference standards, the analysis
conducted in this sheet requires that the results be entered separately. There are
| two ways to do this. One way is to enter an analysis set (one candidate standard | [7 empty] | STEP 6 |  | Assay Candidate Standard on |  |  |  |  | STEP 6 |  | Assay Candidate Standard on | [5 empty] | STEP 6 |  | Assay Candidate Standard on | [25 empty] | STEP 6 |  | Assay Candidate Standard on
| response and the responses from its zero and nonzero standard analyses) | [9 empty] | Different day from Initial Calibration, continued | [6 empty] | Different day from Initial Calibration, continued | [7 empty] | Different day from Initial Calibration, continued | [27 empty] | Different day from Initial Calibration, continued
| below. Another approach is to enter all of the responses (multiple sets) below. | [7 empty] | Calculations - Linear Model | [6 empty] | Calculations - Quadratic Model | [7 empty] | Calculations - Cubic Model | [27 empty] | Calculations - Quartic Model
Enter zero and reference standard responses in the area to the left and enter the
| responses to a single candidate standard in the area to the right, below. | [7 empty] | Var-Covar | [6 empty] | Var-Covar | [7 empty] | Var-Covar | [27 empty] | Var-Covar
[9 empty] | 0.13657 | -0.022147 | 0 | 0 |  |  |  | 0.00019741 | -8.5262e-05 | 7.6813e-06 | 0 | 0 | 0 |  |  | 4560.2545 | -3674.2826 | 793.0877 | -50.0145 | 0 | 0 | 0 | 0 | [20 empty] | 355.8361 | -453.1953 | 170.3207 | -24.497 | 1.1917 | 0 | 0 | 0 | 0 | 0
| STEP 6 | (Continued) | [6 empty] | -0.022147 | 0.0049215 | 0 | 0 |  |  |  | -8.5262e-05 | 5.2383e-05 | -5.4221e-06 | 0 | 0 | 0 |  |  | -3674.2826 | 4460.7391 | -1126.9033 | 77.3101 | 0 | 0 | 0 | 0 | [20 empty] | -453.1953 | 945.3718 | -426.9248 | 67.6951 | -3.4993 | 0 | 0 | 0 | 0 | 0
[9 empty] | 0 | 0 | 0.19481 | -0.025974 |  |  |  | 7.6813e-06 | -5.4221e-06 | 6.0245e-07 | 0 | 0 | 0 |  |  | 793.0877 | -1126.9033 | 307.1271 | -22.0652 | 0 | 0 | 0 | 0 | [20 empty] | 170.3207 | -426.9248 | 209.2522 | -34.8411 | 1.861 | 0 | 0 | 0 | 0 | 0
| See table in Step 6 of the Instructions for the minimum number of | [7 empty] | 0 | 0 | -0.025974 | 0.0057721 |  |  |  | 0 | 0 | 0 | 0.00017027 | -5.6756e-05 | 4.2042e-06 |  |  | -50.0145 | 77.3101 | -22.0652 | 1.6345 | 0 | 0 | 0 | 0 | [20 empty] | -24.497 | 67.6951 | -34.8411 | 5.9882 | -0.32715 | 0 | 0 | 0 | 0 | 0
| different nonzero reference standard concentrations required. | [14 empty] | 0 | 0 | 0 | -5.6756e-05 | 5.4654e-05 | -5.6055e-06 |  |  | 0 | 0 | 0 | 0 | 2613.2598 | 0 | -6525.7809 | 407.8613 | [20 empty] | 1.1917 | -3.4993 | 1.861 | -0.32715 | 0.018175 | 0 | 0 | 0 | 0 | 0
[9 empty] | Have Coefficients Remained the Same? | [6 empty] | 0 | 0 | 0 | 4.2042e-06 | -5.6055e-06 | 6.2284e-07 |  |  | 0 | 0 | 0 | 0 | 14808.4722 | -5.5857e+18 | 1.8619e+18 | -1.3792e+17 | [20 empty] | 0 | 0 | 0 | 0 | 0 | 159.1315 | 0 | -70.1259 | 0 | 0
| Estimates below are based on the quadratic model. | [7 empty] | b0-b2 | -1.7673 | 0.53535 | TRUE | [11 empty] | 0 | 0 | 0 | 0 | -5161.9947 | 1.8619e+18 | -6.2063e+17 | 45972952624875336 | [20 empty] | 0 | 0 | 0 | 0 | 0 | -128.722 | 80879167504708432 | -13563533334747200 | -2.4684e+15 | 3.3077e+14
| Tab-Right to view estimates based on the other models. | [7 empty] | b1-b3 | -0.20717 | 0.20647 | TRUE |  |  |  | Have Coefficients Remained the Same? | [7 empty] | 0 | 0 | 0 | 0 | 387.1496 | -1.3792e+17 | 45972952624875360 | -3.4054e+15 | [20 empty] | 0 | 0 | 0 | 0 | 0 | 22.9351 | -20312119692963300 | 3.37e+15 | 6.3205e+14 | -8.3969e+13
| Reference Standards | [5 empty] | Candidate Standard | [8 empty] | b0-b3 | -0.022979 | -0.061328 | 0.015371 | TRUE |  |  |  | Have Coefficients Remained the Same? | [27 empty] | 0 | 0 | 0 | 0 | 0 | 0.20478 | -2.1885e+14 | 4.8835e+13 | 2.6346e+12 | -595430688578.8666
| (Enter 0 for zero Conc.) | [5 empty] | Observed | Estimated | Coefficients are not significantly different. | [6 empty] | b1-b4 | 0.0084949 | -0.012197 | 0.029187 | TRUE |  |  |  | b0-b4 | -0.024229 | -169.4174 | 169.369 | TRUE | [23 empty] | 0 | 0 | 0 | 0 | 0 | -0.15358 | 1.6414e+14 | -2.8425e+13 | -4.7098e+12 | 649081016501.04
| Conc. | Response | Conc.^2 | Conc.^3 | Conc.^4 |  | Response | Conc. | Consider including the new data as part of original calibration (Step 1). | [6 empty] | b2-b5 | -0.00098256 | -0.0031964 | 0.0012313 | TRUE |  |  |  | b1-b5 | -126.997 | #NUM! | #NUM! | #NUM!
| 0 | 0.2184 | 0 | 0 | 0 |  | 4.01 | 2.94 | [15 empty] | b2-b6 | 32.0994 | #NUM! | #NUM! | #NUM! | [23 empty] | Have Coefficients Remained the Same?
| 0 | 0.2194 | 0 | 0 | 0 |  | 4 | 2.9337 | Mean Response |  | 95% Uncert Limits |  |  |  |  | Coefficients are not significantly different. | [7 empty] | b3-b7 | 2 | #NUM! | #NUM! | #NUM! | [23 empty] | b0-b5 | 0.0011643 | -45.3846 | 45.387 | TRUE
| 0 | 0.2204 | 0 | 0 | 0 |  | 3.99 | 2.9274 | 4 |  | 2.8583 | 5.1417 |  |  |  | Consider including the new data as part of original calibration (Step 1). | [35 empty] | b1-b6 | -15.0715 | -568785273.3359 | 568785243.1928 | TRUE
| 4.5 | 6.693 | 20.25 | 91.125 | 410.0625 | [18 empty] | #NUM! | [27 empty] | b2-b7 | 4.139 | -116102800.5247 | 116102808.8028 | TRUE
| 4.5 | 6.723 | 20.25 | 91.125 | 410.0625 |  |  |  | Std Variance of Estimate |  |  | Estimation/Prediction Variance |  |  |  | Mean Response |  | 95% Uncert Limits | [5 empty] | #NUM! | [27 empty] | b3-b8 | 0.49307 | -3246279.1661 | 3246280.1523 | TRUE
| 4.5 | 6.773 | 20.25 | 91.125 | 410.0625 |  |  |  | 0.092083 |  |  | 0.39391 |  |  |  | 4 |  | 3.9646 | 4.0354 | [32 empty] | b4-b9 | -0.062113 | -1611311.3499 | 1611311.2256 | TRUE
| 9 | 17.3171 | 81 | 729 | 6561 | [18 empty] | Mean Response |  | 95% Uncert Limits | [7 empty] | p | q | r | a | b | m | q | Solution 1 | Solution 2 | Solution 3 |  |  |  | Best Solution | Indicator | Unique? | A | B
| 9 | 17.3271 | 81 | 729 | 6561 |  |  |  | Conc. | 95% Uncert Limits | [5 empty] | Std Variance of Estimate |  |  | Estimation/Prediction Variance |  |  |  |  | 4 |  | #NUM! | #NUM! | [6 empty] | 16 | -64 | 1.8903 | -149.3333 | 646.631 | 14.1107 | 0.91347 | 3.2883 | -19.318 | 0.029758 | 3.2883 | 999999 | 0.029758 | 0.029758 | -18807.8041 | #NUM! | #NUM! | #NUM!
| 9 | 17.3371 | 81 | 729 | 6561 |  |  |  | 2.3473 | 1.7466 | 2.9479 |  |  |  |  | 0.00014308 |  |  | 0.45661 | [14 empty] | 16 | -64 | #NUM! | -149.3333 | #NUM! | 14.1107 | #NUM! | #NUM! | #NUM! | #NUM! | #NUM! | #NUM! | #NUM! | #NUM! | #NUM! | #NUM! | #NUM! | #NUM! | Coefficients are not significantly different.
[10 empty] | -25.59% | 25.59% | [12 empty] | Std Variance of Estimate |  |  | Estimation/Prediction Variance | [6 empty] | 16 | -64 | #NUM! | -149.3333 | #NUM! | 14.1107 | #NUM! | #NUM! | #NUM! | #NUM! | #NUM! | #NUM! | #NUM! | #NUM! | #NUM! | #NUM! | #NUM! | #NUM! | Consider including the new data as part of original calibration (Step 1).
[7 empty] | nnn = | 3 | [7 empty] | Conc. | 95% Uncert Limits | [6 empty] | -4848849335615375 |  |  | 1
[6 empty] | mean = | 4 | 2.9337 | [7 empty] | 2.9337 | 2.9114 | 2.956 | [33 empty] | Mean Response |  | 95% Uncert Limits
[7 empty] | Experim. std. dev. = | 0.0063102 | [7 empty] | 0.7615% | -0.7626% | 0.7604% | [5 empty] | Conc. | 95% Uncert Limits | [26 empty] | 4 |  | -530422512.7709 | 530422520.7709
[7 empty] | Rel. Std. Uncertainty = | 0.1242% | [15 empty] | 0.029758 | #NUM! | #NUM! | [37 empty] | quartic equation coefficients (r.h.s.=0) |  |  |  | resolvent cubic equationcoefficients
[7 empty] | df = | 2 | [15 empty] | #NUM! | #NUM! | #NUM! | [25 empty] | Std Variance of Estimate |  |  | Estimation/Prediction Variance | [8 empty] | a | b | c | d | pp | qq | rr | aa | bb | Indicator | mm | q | Conc1 | Conc2 | Conc3 | A | B | yy | R | D | E | Conc1 | Conc2 | Conc3 | Conc4 | In calibration range? |  |  |  | Best
[6 empty] | F = | 0.077954 | [16 empty] | Portion of Uncertainty2 due to calibration uncertainty |  |  |  | 100% | [23 empty] | 70337011574383072 |  |  | 1 | [8 empty] | -17.8945 | 359.024 | 2397.5035 | -9759.3755 | -359.024 | -3864.5976 | -16638348.5675 | -46830.6836 | -20528811.39 | 1.0155e+14 | 249.8818 | #NUM! | #NUM! | #NUM! | #NUM! | 272.9794 | 57.1846 | 449.8387 | 13.0716 | #NUM! | 9.9984 | #NUM! | #NUM! | 2.937 | -7.0614 | 0 | 0 | 2.937 | 0 | 2.937
[6 empty] | F sig? = | FALSE | [56 empty] | -17.8945 | 359.024 | 2397.5035 | 1.3697e+12 | -359.024 | -5.4787e+12 | 1.5284e+15 | -5.4787e+12 | 8.7273e+14 | -6.0907e+36 | 2702762.4539 | 0.52366 | 2340700.968 | -2340620.9148 | 278.9709 | #NUM! | #NUM! | 2340700.968 | 1529.8438 | #NUM! | #NUM! | #NUM! | #NUM! | #NUM! | #NUM! | 0 | 0 | 0 | 0 | 0
[6 empty] | Pr{>F} = | 0.92526 | [44 empty] | Conc. | 95% Uncert Limits | [10 empty] | -17.8945 | 359.024 | 2397.5035 | -1.3697e+12 | -359.024 | 5.4787e+12 | -1.5284e+15 | 5.4787e+12 | -8.7273e+14 | 6.0907e+36 | #NUM! | #NUM! | #NUM! | #NUM! | #NUM! | 1351460.8658 | -1351301.5696 | 278.9709 | 0.0020906 | #NUM! | 2163.5497 | #NUM! | #NUM! | 1086.2474 | -1077.3023 | 0 | 0 | 0 | 0 | 0
[52 empty] | 2.937 | 0 | 0
[6 empty] | The relative standard uncertainty is okay. | [45 empty] | 0% | -100% | -100%
| nn = | 9 | [49 empty] | Portion of Uncertainty2 due to calibration uncertainty |  |  |  | 100%
Coefficients are not significantly different.
Consider including the new data as part of original calibration (Step 1).
Estimated Concentration of Candidate Standard
2.9337
| Rel. Expanded Uncertainty |  | Portion of Variance Due to Calibration Uncertainty
| 0.7615% |  | 45.66%
Expanded Uncertainty for Candidate Standard Concentration
| Lower | Upper |  | The approximate level of confidence is 95%.
| 2.9114 | 2.956
These upper and lower limits are compared with the corresponding
limits estimated on different assay dates to establish that the
candidate standard has not drifted.
Note on U53: This compares the residuals of the linear and quadratic regressions. If significant, then quadratic is the preferred model.
Note on AD53: This compares the residuals of the quadratic and cubic regressions. If significant, then cubic is the preferred model.
Note on BG53: This compares the residuals of the quadratic and cubic regressions. If significant, then cubic is the preferred model.
Note on U56: Comparing the Linear and Quadratic solutions. Q --> quadratic is better L --> linear is better model
Note on AD56: Comparing the Quadratic and Cubic solutions
Note on BG56: Comparing the Cubic and Quartic solutions
Note on AC177: The Lower and Upper limits shown here assume linearity of response in the neighborhood of this concentration.
Note on AD177: The Lower and Upper limits shown here assume linearity of response in the neighborhood of this concentration.
Note on BF177: Mike Messner: The Lower and Upper limits shown here assume linearity of response in the neighborhood of this concentration.
Note on BG177: Mike Messner: The Lower and Upper limits shown here assume linearity of response in the neighborhood of this concentration.
Note on S182: #NUM indicates that the response is out of range. The quadratic curve has either a maximum or minimum at -b1/(2*b2), depending on whether b2 is negative or positive, respectively.
Note on BE183: Zero here indicates that no concentrations in the calibration range were expected to produce the observed response.
Note on BR183: x^4+a*x^3+b*x^2+c*x+d=0
Note on BV183: yy^3+pp*yy^2+qq*yy+rr=0
Note on AL184: b2/b3
Note on AM184: b1/b3
Note on AN184: (b0-response)/b3
Note on AO184: (3*q-p^2)/3
Note on AP184: (2*p^3-9*p*q+27*r)/27
Note on AQ184: 2*sqrt(-a/3)
Note on AR184: acos(3*b/(a*m))/3
Note on AS184: 2*sqrt(-a/3)*cos(q)-p/3
Note on AT184: 2*sqrt(-a/3)*cos(q+2*pi/3))-p/3
Note on AU184: 2*sqrt(-a/3)*cos(q+4*pi/3))-p/3
Note on BR185: b3/b4
Note on BS185: b2/b4
Note on BT185: b1/b4
Note on BU185: (b0-response)/b4
Note on BV185: -b
Note on BW185: a*c-4*d
Note on BX185: a^2*d-4*b*d-c^2
Note on BY185: (3*qq-pp^2)/3
Note on BZ185: (3*pp^3-9*pp*qq+27*rr)/27
Note on CA185: bb^2/4+aa^3/27 If positive, the one real root is given by A+B-p/3. If negative, there will be three real roots. Hopefully, only one will be in the calibration range!
Note on CB185: 2*sqrt(-aa/3) Will be imaginary if aa>0.
Note on CC185: acos(3*bb/(aa*mm)) May be complex.
Note on CD185: mm*cos(theta)-p/3 May be complex.
Note on CE185: mm*cos(theta+2pi/3)-p/3 May be complex.
Note on CF185: mm*cos(theta+4pi/3)-p/3 May be complex.
Note on CG185: (-b/2+sqrt(indicator))^(1/3) Used if the indicator is positive.
Note on CH185: (-b/2-sqrt(indicator))^(1/3)-p/3 Used if the indicator is positive.
Note on CI185: =largest of the three solutioins if indicator<0. =A+B-p/3 if indicator>0.
Note on CJ185: sqrt(a^2/4-b+yy)
Note on CK185: =sqrt(3*a^2/4-R^2+(4*a*b-8*c-a^3)/(4*R)) Note: R cannot be zero
Note on CL185: =sqrt(3*a^2/4-R^2-(4*a*b-8*c-a^3)/(4*R)) Note: R cannot be zero
Note on CM185: =-a/4+R/2+D/2 May be complex.
Note on CN185: =-a/4+R/2-D/2 May be complex.
Note on CO185: =-a/4-R/2+E/2 May be complex.
Note on CP185: =-a/4-R/2-E/2 May be complex.
Note on CQ185: Returns the concentration, if in the calibration range…zero otherwise.
Note on CR185: Returns the concentration, if in the calibration range…zero otherwise.
Note on CS185: Returns the concentration, if in the calibration range…zero otherwise.
Note on CT185: Returns the concentration, if in the calibration range…zero otherwise.
Note on CU185: Returns a solution, if unique. Returns an error message if multiple solutions exist in the calibration range.
Note on AL216: b6/b7
Note on AM216: b5/b7
Note on AN216: (b4-response)/b7
Note on AO216: (3*q-p^2)/3
Note on AP216: (2*p^3-9*p*q+27*r)/27
Note on AQ216: 2*sqrt(-a/3)
Note on AR216: acos(3*b/(a*m))/3
Note on AS216: 2*sqrt(-a/3)*cos(q)-p/3
Note on AT216: 2*sqrt(-a/3)*cos(q+2*pi/3))-p/3
Note on AU216: 2*sqrt(-a/3)*cos(q+4*pi/3))-p/3
Note on AE225: The Lower and Upper limits shown here assume linearity of response in the neighborhood of this concentration.
Note on AF225: The Lower and Upper limits shown here assume linearity of response in the neighborhood of this concentration.
Note on BH225: Mike Messner: The Lower and Upper limits shown here assume linearity of response in the neighborhood of this concentration.
Note on BI225: Mike Messner: The Lower and Upper limits shown here assume linearity of response in the neighborhood of this concentration.
Note on L231: If "TRUE" then the observed responses varied more than expected.
Note on T231: If "TRUE" then the observed responses varied more than expected.
Note on AC231: If "TRUE" then the observed responses varied more than expected.
Note on BF231: : If "TRUE" then the observed responses varied more than expected.
Note on AT266: 2*sqrt(-a/3)*cos(q+2*pi/3))-p/3
Note on Z276: This estimate is derived from the mean response. This is not the same as the mean of the above estimates.
Note on AK277: This estimate is derived from the mean response. This is not the same as the mean of the above estimates.
Note on BP277: This estimate is derived from the mean response. This is not the same as the mean of the above estimates.

## Chart Data
This sheet contains the chart data. The chart data include initial calibration data and the
regression information derived from those data.
Only the initial calibration data (Step 1 in Measurement Data worksheet) are included.
| Recalibration data (Step 6) are not included. | [5 empty] | 0.95 | = Confidence Level for Plotting Confidence Bands
0 | = minimum concentration | [5 empty] | 2.1098 | = t for linear model
9 | = maximum concentration | [5 empty] | 2.1199 | = t for quadratic model
19 | = n, the number of initial calibration points | [5 empty] | 2.1314 | = t for cubic model
[7 empty] | 2.1448 | = t for quartic model
Initial Calibration Data |  |  |  |  | Estimates, Linear Model |  |  |  | Estimates, Quadratic Model |  |  |  | Estimates, Cubic Model |  |  |  | Estimates, Quartic Model
i | Xi | Yi |  |  | Estimate | Lower | Upper | Residual | Estimate | Lower | Upper | Residual | Estimate | Lower | Upper | Residual | Estimate | Lower | Upper | Residual
1 | 0 | 0.2194 | [5 empty] | 1.2972 |  |  |  | 0.022979 |  |  |  | 0.024229 |  |  |  | -0.0011643
2 | 0.5 | 0.7141 | [5 empty] | 0.8416 |  |  |  | -0.0078507 |  |  |  | -0.0074339 |  |  |  | -0.0046125
3 | 1 | 1.2885 | [5 empty] | 0.46574 |  |  |  | -0.0089477 |  |  |  | -0.0091193 |  |  |  | 0.0069798
4 | 1.5 | 1.9132 | [5 empty] | 0.14019 |  |  |  | -0.0097122 |  |  |  | -0.010258 |  |  |  | 0.0085384
5 | 2 | 2.591 | [5 empty] | -0.13226 |  |  |  | -0.0073442 |  |  |  | -0.0080797 |  |  |  | 0.0066086
6 | 2.5 | 3.2866 | [5 empty] | -0.38692 |  |  |  | -0.037144 |  |  |  | -0.037916 |  |  |  | -0.030945
7 | 3 | 4.1078 | [5 empty] | -0.51597 |  |  |  | 0.0086889 |  |  |  | 0.0080025 |  |  |  | 0.0062598
8 | 3.5 | 4.9446 | [5 empty] | -0.62942 |  |  |  | 0.020154 |  |  |  | 0.019645 |  |  |  | 0.010227
9 | 4 | 5.8145 | [5 empty] | -0.70978 |  |  |  | 0.014752 |  |  |  | 0.014482 |  |  |  | -0.00012318
10 | 4.5 | 6.723 | [5 empty] | -0.75153 |  |  |  | -0.0020181 |  |  |  | -0.0020181 |  |  |  | -0.018449
11 | 5 | 7.7284 | [5 empty] | -0.69639 |  |  |  | 0.028144 |  |  |  | 0.028414 |  |  |  | 0.013809
12 | 5.5 | 8.7566 | [5 empty] | -0.61844 |  |  |  | 0.031139 |  |  |  | 0.031648 |  |  |  | 0.022229
13 | 6 | 9.8013 | [5 empty] | -0.52399 |  |  |  | 0.00066681 |  |  |  | 0.0013533 |  |  |  | -0.00038941
14 | 6.5 | 10.8818 | [5 empty] | -0.39375 |  |  |  | -0.043973 |  |  |  | -0.043201 |  |  |  | -0.03623
15 | 7 | 12.0982 | [5 empty] | -0.1276 |  |  |  | -0.0026811 |  |  |  | -0.0019456 |  |  |  | 0.012743
16 | 7.5 | 13.3122 | [5 empty] | 0.13615 |  |  |  | -0.013756 |  |  |  | -0.013211 |  |  |  | 0.0055851
17 | 8 | 14.584 | [5 empty] | 0.45769 |  |  |  | -0.016999 |  |  |  | -0.016828 |  |  |  | -0.00072859
18 | 8.5 | 15.9238 | [5 empty] | 0.84724 |  |  |  | -0.0022097 |  |  |  | -0.0026265 |  |  |  | 0.00019503
19 | 9 | 17.3271 | [5 empty] | 1.3003 |  |  |  | 0.026112 |  |  |  | 0.024862 |  |  |  | -0.00053133
20 | -999 | -999 | [5 empty] | 1.3003 |  |  |  | 0.026112 |  |  |  | 0.024862 |  |  |  | -0.00053133
21 | -999 | -999 | [5 empty] | 1.3003 |  |  |  | 0.026112 |  |  |  | 0.024862 |  |  |  | -0.00053133
22 | -999 | -999 | [5 empty] | 1.3003 |  |  |  | 0.026112 |  |  |  | 0.024862 |  |  |  | -0.00053133
23 | -999 | -999 | [5 empty] | 1.3003 |  |  |  | 0.026112 |  |  |  | 0.024862 |  |  |  | -0.00053133
24 | -999 | -999 | [5 empty] | 1.3003 |  |  |  | 0.026112 |  |  |  | 0.024862 |  |  |  | -0.00053133
25 | -999 | -999 | [5 empty] | 1.3003 |  |  |  | 0.026112 |  |  |  | 0.024862 |  |  |  | -0.00053133
26 | -999 | -999 | [5 empty] | 1.3003 |  |  |  | 0.026112 |  |  |  | 0.024862 |  |  |  | -0.00053133
27 | -999 | -999 | [5 empty] | 1.3003 |  |  |  | 0.026112 |  |  |  | 0.024862 |  |  |  | -0.00053133
28 | -999 | -999 | [5 empty] | 1.3003 |  |  |  | 0.026112 |  |  |  | 0.024862 |  |  |  | -0.00053133
29 | -999 | -999 | [5 empty] | 1.3003 |  |  |  | 0.026112 |  |  |  | 0.024862 |  |  |  | -0.00053133
30 | -999 | -999 | [5 empty] | 1.3003 |  |  |  | 0.026112 |  |  |  | 0.024862 |  |  |  | -0.00053133
31 | -999 | -999 | [5 empty] | 1.3003 |  |  |  | 0.026112 |  |  |  | 0.024862 |  |  |  | -0.00053133
32 | -999 | -999 | [5 empty] | 1.3003 |  |  |  | 0.026112 |  |  |  | 0.024862 |  |  |  | -0.00053133
33 | -999 | -999 | [5 empty] | 1.3003 |  |  |  | 0.026112 |  |  |  | 0.024862 |  |  |  | -0.00053133
34 | -999 | -999 | [5 empty] | 1.3003 |  |  |  | 0.026112 |  |  |  | 0.024862 |  |  |  | -0.00053133
35 | -999 | -999 | [5 empty] | 1.3003 |  |  |  | 0.026112 |  |  |  | 0.024862 |  |  |  | -0.00053133
36 | -999 | -999 | [5 empty] | 1.3003 |  |  |  | 0.026112 |  |  |  | 0.024862 |  |  |  | -0.00053133
37 | -999 | -999 | [5 empty] | 1.3003 |  |  |  | 0.026112 |  |  |  | 0.024862 |  |  |  | -0.00053133
38 | -999 | -999 | [5 empty] | 1.3003 |  |  |  | 0.026112 |  |  |  | 0.024862 |  |  |  | -0.00053133
39 | -999 | -999 | [5 empty] | 1.3003 |  |  |  | 0.026112 |  |  |  | 0.024862 |  |  |  | -0.00053133
40 | -999 | -999 | [5 empty] | 1.3003 |  |  |  | 0.026112 |  |  |  | 0.024862 |  |  |  | -0.00053133
41 | -999 | -999 | [5 empty] | 1.3003 |  |  |  | 0.026112 |  |  |  | 0.024862 |  |  |  | -0.00053133
42 | -999 | -999 | [5 empty] | 1.3003 |  |  |  | 0.026112 |  |  |  | 0.024862 |  |  |  | -0.00053133
43 | -999 | -999 | [5 empty] | 1.3003 |  |  |  | 0.026112 |  |  |  | 0.024862 |  |  |  | -0.00053133
44 | -999 | -999 | [5 empty] | 1.3003 |  |  |  | 0.026112 |  |  |  | 0.024862 |  |  |  | -0.00053133
45 | -999 | -999 | [5 empty] | 1.3003 |  |  |  | 0.026112 |  |  |  | 0.024862 |  |  |  | -0.00053133
46 | -999 | -999 | [5 empty] | 1.3003 |  |  |  | 0.026112 |  |  |  | 0.024862 |  |  |  | -0.00053133
47 | -999 | -999 | [5 empty] | 1.3003 |  |  |  | 0.026112 |  |  |  | 0.024862 |  |  |  | -0.00053133
48 | -999 | -999 | [5 empty] | 1.3003 |  |  |  | 0.026112 |  |  |  | 0.024862 |  |  |  | -0.00053133
49 | -999 | -999 | [5 empty] | 1.3003 |  |  |  | 0.026112 |  |  |  | 0.024862 |  |  |  | -0.00053133
50 | -999 | -999 | [5 empty] | 1.3003 |  |  |  | 0.026112 |  |  |  | 0.024862 |  |  |  | -0.00053133
[21 empty] | x'V for Linear |  | x'V for Quadratic |  |  | x'V for Cubic |  |  |  | x'V for Quartic
1 | 0 | 0 | 0 | 0 | -1.0778 | -1.7351 | -0.42036 |  | 0.19642 | 0.16738 | 0.22546 |  | 0.19517 | 0.17064 | 0.21971 |  | 0.22056 | 0.19967 | 0.24145 |  | 0.097088 | -0.015744 | 0.00018765 | -8.1049e-05 | 7.3017e-06 | 0.00028244 | -0.00022757 | 4.9121e-05 | -3.0977e-06 | 0.00020347 | -0.00025914 | 9.7392e-05 | -1.4008e-05 | 6.8146e-07
1 | 0.18 | 0.0324 | 0.005832 | 0.0010498 | -0.73566 | -1.374 | -0.097346 |  | 0.37986 | 0.353 | 0.40671 |  | 0.37894 | 0.35769 | 0.40018 |  | 0.39215 | 0.37538 | 0.40893 |  | 0.094254 | -0.015114 | 0.0001733 | -7.2253e-05 | 6.3925e-06 | 0.00024305 | -0.00018007 | 3.7166e-05 | -2.2795e-06 | 0.0001599 | -0.00016953 | 5.7212e-05 | -7.6659e-06 | 3.5469e-07
1 | 0.36 | 0.1296 | 0.046656 | 0.016796 | -0.39357 | -1.013 | 0.22589 |  | 0.56977 | 0.54494 | 0.59459 |  | 0.56914 | 0.55063 | 0.58765 |  | 0.57251 | 0.55853 | 0.5865 |  | 0.09142 | -0.014484 | 0.00015942 | -6.3791e-05 | 5.5204e-06 | 0.00020674 | -0.00013693 | 2.6396e-05 | -1.5463e-06 | 0.00012216 | -9.4402e-05 | 2.4103e-05 | -2.4979e-06 | 9.0475e-08
1 | 0.54 | 0.2916 | 0.15746 | 0.085031 | -0.051478 | -0.65233 | 0.54937 |  | 0.76615 | 0.74319 | 0.78911 |  | 0.76579 | 0.74945 | 0.78213 |  | 0.76142 | 0.74902 | 0.77382 |  | 0.088586 | -0.013855 | 0.00014602 | -5.5663e-05 | 4.6855e-06 | 0.00017339 | -9.7977e-05 | 1.6763e-05 | -8.946e-07 | 8.9786e-05 | -3.2493e-05 | -2.5896e-06 | 1.6089e-06 | -1.1733e-07
1 | 0.72 | 0.5184 | 0.37325 | 0.26874 | 0.29061 | -0.2919 | 0.87313 |  | 0.96901 | 0.94775 | 0.99028 |  | 0.96888 | 0.95416 | 0.9836 |  | 0.95865 | 0.9469 | 0.9704 |  | 0.085752 | -0.013225 | 0.00013308 | -4.7869e-05 | 3.8876e-06 | 0.0001429 | -6.3043e-05 | 8.2187e-06 | -3.2082e-07 | 6.2332e-05 | 1.743e-05 | -2.3498e-05 | 4.7626e-06 | -2.7461e-07
1 | 0.9 | 0.81 | 0.729 | 0.6561 | 0.63271 | 0.068225 | 1.1972 |  | 1.1784 | 1.1586 | 1.1981 |  | 1.1784 | 1.1648 | 1.192 |  | 1.164 | 1.1523 | 1.1757 |  | 0.082918 | -0.012595 | 0.00012062 | -4.0409e-05 | 3.1269e-06 | 0.00011516 | -3.1962e-05 | 7.1612e-07 | 1.7857e-07 | 3.9366e-05 | 5.6543e-05 | -3.9224e-05 | 7.0665e-06 | -3.87e-07
1 | 1.08 | 1.1664 | 1.2597 | 1.3605 | 0.9748 | 0.42803 | 1.5216 |  | 1.3942 | 1.3757 | 1.4126 |  | 1.3944 | 1.3814 | 1.4074 |  | 1.3773 | 1.3655 | 1.3891 |  | 0.080085 | -0.011965 | 0.00010864 | -3.3282e-05 | 2.4032e-06 | 9.006e-05 | -4.5657e-06 | -5.7929e-06 | 6.0712e-07 | 2.0476e-05 | 8.5976e-05 | -5.0345e-05 | 8.6191e-06 | -4.5988e-07
1 | 1.26 | 1.5876 | 2.0004 | 2.5205 | 1.3169 | 0.78746 | 1.8463 |  | 1.6165 | 1.5992 | 1.6337 |  | 1.6168 | 1.6041 | 1.6295 |  | 1.5983 | 1.5863 | 1.6102 |  | 0.077251 | -0.011336 | 9.7125e-05 | -2.649e-05 | 1.7167e-06 | 6.7493e-05 | 1.9313e-05 | -1.1356e-05 | 9.6836e-07 | 5.2673e-06 | 0.00010681 | -5.741e-05 | 9.5145e-06 | -4.9834e-07
1 | 1.44 | 2.0736 | 2.986 | 4.2998 | 1.659 | 1.1465 | 2.1715 |  | 1.8452 | 1.8289 | 1.8616 |  | 1.8457 | 1.8331 | 1.8584 |  | 1.8268 | 1.8148 | 1.8388 |  | 0.074417 | -0.010706 | 8.6084e-05 | -2.0032e-05 | 1.0673e-06 | 4.735e-05 | 3.9842e-05 | -1.6022e-05 | 1.2658e-06 | -6.6397e-06 | 0.00012006 | -6.0943e-05 | 9.8417e-06 | -5.0722e-07
1 | 1.62 | 2.6244 | 4.2515 | 6.8875 | 2.0011 | 1.5051 | 2.497 |  | 2.0805 | 2.0649 | 2.0961 |  | 2.0811 | 2.0683 | 2.0938 |  | 2.0628 | 2.0508 | 2.0747 |  | 0.071583 | -0.010076 | 7.5517e-05 | -1.3908e-05 | 4.5494e-07 | 2.9522e-05 | 5.7189e-05 | -1.9837e-05 | 1.5031e-06 | -1.5606e-05 | 0.00012671 | -6.1439e-05 | 9.6852e-06 | -4.9112e-07
1 | 1.8 | 3.24 | 5.832 | 10.4976 | 2.3432 | 1.8632 | 2.8231 |  | 2.3222 | 2.3071 | 2.3372 |  | 2.3229 | 2.3099 | 2.3358 |  | 2.3059 | 2.2942 | 2.3176 |  | 0.068749 | -0.0094464 | 6.5423e-05 | -8.1178e-06 | -1.2026e-07 | 1.3902e-05 | 7.1521e-05 | -2.285e-05 | 1.6837e-06 | -2.1976e-05 | 0.00012769 | -5.9368e-05 | 9.1248e-06 | -4.5435e-07
1 | 1.98 | 3.9204 | 7.7624 | 15.3695 | 2.6853 | 2.2208 | 3.1497 |  | 2.5704 | 2.5557 | 2.585 |  | 2.5711 | 2.558 | 2.5842 |  | 2.5562 | 2.5448 | 2.5676 |  | 0.065915 | -0.0088166 | 5.5803e-05 | -2.6616e-06 | -6.5836e-07 | 3.8175e-07 | 8.3005e-05 | -2.5109e-05 | 1.8111e-06 | -2.6077e-05 | 0.00012387 | -5.5172e-05 | 8.2356e-06 | -4.0097e-07
1 | 2.16 | 4.6656 | 10.0777 | 21.7678 | 3.0273 | 2.5778 | 3.4769 |  | 2.825 | 2.8106 | 2.8395 |  | 2.8258 | 2.8126 | 2.839 |  | 2.8133 | 2.8023 | 2.8244 |  | 0.063081 | -0.0081869 | 4.6655e-05 | 2.4606e-06 | -1.1593e-06 | -1.1148e-05 | 9.181e-05 | -2.6661e-05 | 1.889e-06 | -2.8219e-05 | 0.00011607 | -4.9267e-05 | 7.0878e-06 | -3.3477e-07
1 | 2.34 | 5.4756 | 12.8129 | 29.9822 | 3.3694 | 2.9342 | 3.8047 |  | 3.0862 | 3.0718 | 3.1006 |  | 3.087 | 3.0737 | 3.1002 |  | 3.0773 | 3.0666 | 3.0879 |  | 0.060247 | -0.0075571 | 3.7981e-05 | 7.2488e-06 | -1.6232e-06 | -2.0796e-05 | 9.8103e-05 | -2.7554e-05 | 1.9208e-06 | -2.8694e-05 | 0.00010508 | -4.2041e-05 | 5.7472e-06 | -2.5929e-07
1 | 2.52 | 6.3504 | 16.003 | 40.3276 | 3.7115 | 3.2898 | 4.1333 |  | 3.3538 | 3.3394 | 3.3682 |  | 3.3546 | 3.3413 | 3.3678 |  | 3.3479 | 3.3377 | 3.3582 |  | 0.057413 | -0.0069274 | 2.9779e-05 | 1.1703e-05 | -2.05e-06 | -2.867e-05 | 0.00010205 | -2.7836e-05 | 1.9101e-06 | -2.7778e-05 | 9.1608e-05 | -3.3856e-05 | 4.2745e-06 | -1.7781e-07
1 | 2.7 | 7.29 | 19.683 | 53.1441 | 4.0536 | 3.6446 | 4.4626 |  | 3.6279 | 3.6134 | 3.6424 |  | 3.6286 | 3.6155 | 3.6418 |  | 3.6252 | 3.6152 | 3.6352 |  | 0.054579 | -0.0062976 | 2.2051e-05 | 1.5823e-05 | -2.4396e-06 | -3.4878e-05 | 0.00010382 | -2.7556e-05 | 1.8604e-06 | -2.5729e-05 | 7.6339e-05 | -2.5049e-05 | 2.726e-06 | -9.3336e-08
1 | 2.88 | 8.2944 | 23.8879 | 68.7971 | 4.3957 | 3.9986 | 4.7928 |  | 3.9085 | 3.8938 | 3.9232 |  | 3.9092 | 3.8962 | 3.9222 |  | 3.9089 | 3.8991 | 3.9186 |  | 0.051745 | -0.0056678 | 1.4796e-05 | 1.961e-05 | -2.7922e-06 | -3.953e-05 | 0.00010358 | -2.6759e-05 | 1.7754e-06 | -2.2788e-05 | 5.9892e-05 | -1.5926e-05 | 1.1531e-06 | -8.6128e-09
1 | 3.06 | 9.3636 | 28.6526 | 87.677 | 4.7378 | 4.3517 | 5.1239 |  | 4.1955 | 4.1806 | 4.2104 |  | 4.1962 | 4.1834 | 4.209 |  | 4.1989 | 4.1893 | 4.2086 |  | 0.048911 | -0.0050381 | 8.0144e-06 | 2.3062e-05 | -3.1076e-06 | -4.2732e-05 | 0.0001015 | -2.5496e-05 | 1.6584e-06 | -1.918e-05 | 4.2842e-05 | -6.7699e-06 | -3.9731e-07 | 7.3869e-08
1 | 3.24 | 10.4976 | 34.0122 | 110.1996 | 5.0799 | 4.7037 | 5.4561 |  | 4.489 | 4.4739 | 4.5042 |  | 4.4896 | 4.4771 | 4.5022 |  | 4.4953 | 4.4856 | 4.505 |  | 0.046077 | -0.0044083 | 1.7057e-06 | 2.618e-05 | -3.3859e-06 | -4.4594e-05 | 9.7749e-05 | -2.3812e-05 | 1.5131e-06 | -1.5112e-05 | 2.5711e-05 | 2.1646e-06 | -1.8833e-06 | 1.5188e-07
1 | 3.42 | 11.6964 | 40.0017 | 136.8058 | 5.422 | 5.0546 | 5.7894 |  | 4.789 | 4.7737 | 4.8044 |  | 4.7896 | 4.7773 | 4.8018 |  | 4.7979 | 4.7882 | 4.8077 |  | 0.043244 | -0.0037786 | -4.1299e-06 | 2.8964e-05 | -3.6272e-06 | -4.5225e-05 | 9.2487e-05 | -2.1757e-05 | 1.343e-06 | -1.0772e-05 | 8.9701e-06 | 1.065e-05 | -3.2674e-06 | 2.2347e-07
1 | 3.6 | 12.96 | 46.656 | 167.9616 | 5.7641 | 5.4043 | 6.1238 |  | 5.0955 | 5.0799 | 5.1111 |  | 5.096 | 5.084 | 5.1079 |  | 5.1067 | 5.0968 | 5.1166 |  | 0.04041 | -0.0031488 | -9.4922e-06 | 3.1415e-05 | -3.8313e-06 | -4.4732e-05 | 8.5886e-05 | -1.9378e-05 | 1.1516e-06 | -6.3342e-06 | -6.9575e-06 | 1.8484e-05 | -4.5171e-06 | 2.8692e-07
1 | 3.78 | 14.2884 | 54.0102 | 204.1584 | 6.1062 | 5.7528 | 6.4595 |  | 5.4085 | 5.3927 | 5.4242 |  | 5.4088 | 5.3972 | 5.4205 |  | 5.4215 | 5.4115 | 5.4316 |  | 0.037576 | -0.002519 | -1.4381e-05 | 3.3531e-05 | -3.9983e-06 | -4.3223e-05 | 7.8113e-05 | -1.6723e-05 | 9.4255e-07 | -1.9534e-06 | -2.1701e-05 | 2.5494e-05 | -5.6044e-06 | 3.408e-07
1 | 3.96 | 15.6816 | 62.0991 | 245.9126 | 6.4483 | 6.0999 | 6.7966 |  | 5.7279 | 5.712 | 5.7438 |  | 5.7282 | 5.7167 | 5.7396 |  | 5.7425 | 5.7323 | 5.7527 |  | 0.034742 | -0.0018893 | -1.8798e-05 | 3.5313e-05 | -4.1281e-06 | -4.0808e-05 | 6.9336e-05 | -1.3839e-05 | 7.1931e-07 | 2.2322e-06 | -3.4939e-05 | 3.1531e-05 | -6.5062e-06 | 3.8392e-07
1 | 4.14 | 17.1396 | 70.9579 | 293.7659 | 6.7903 | 6.4456 | 7.1351 |  | 6.0538 | 6.0378 | 6.0698 |  | 6.054 | 6.0427 | 6.0652 |  | 6.0695 | 6.0591 | 6.0798 |  | 0.031908 | -0.0012595 | -2.274e-05 | 3.6761e-05 | -4.2209e-06 | -3.7595e-05 | 5.9723e-05 | -1.0774e-05 | 4.8544e-07 | 6.1016e-06 | -4.64e-05 | 3.6475e-05 | -7.204e-06 | 4.1538e-07
1 | 4.32 | 18.6624 | 80.6216 | 348.2852 | 7.1324 | 6.7899 | 7.4749 |  | 6.3862 | 6.3701 | 6.4023 |  | 6.3863 | 6.3751 | 6.3974 |  | 6.4025 | 6.392 | 6.4129 |  | 0.029074 | -0.00062976 | -2.621e-05 | 3.7876e-05 | -4.2766e-06 | -3.3692e-05 | 4.944e-05 | -7.5772e-06 | 2.4449e-07 | 9.5509e-06 | -5.5866e-05 | 4.0232e-05 | -7.6839e-06 | 4.3452e-07
1 | 4.5 | 20.25 | 91.125 | 410.0625 | 7.4745 | 7.1328 | 7.8163 |  | 6.725 | 6.7089 | 6.7411 |  | 6.725 | 6.7139 | 6.7361 |  | 6.7414 | 6.731 | 6.7519 |  | 0.02624 | 6.9389e-18 | -2.9207e-05 | 3.8656e-05 | -4.2951e-06 | -2.9207e-05 | 3.8656e-05 | -4.2951e-06 | -2.3717e-20 | 1.2493e-05 | -6.3165e-05 | 4.2735e-05 | -7.937e-06 | 4.4094e-07
1 | 4.68 | 21.9024 | 102.5032 | 479.7151 | 7.8166 | 7.4741 | 8.1591 |  | 7.0703 | 7.0543 | 7.0864 |  | 7.0702 | 7.0591 | 7.0814 |  | 7.0864 | 7.076 | 7.0969 |  | 0.023406 | 0.00062976 | -3.173e-05 | 3.9103e-05 | -4.2766e-06 | -2.4249e-05 | 2.7538e-05 | -9.7593e-07 | -2.4449e-07 | 1.486e-05 | -6.8179e-05 | 4.3943e-05 | -7.9588e-06 | 4.3452e-07
1 | 4.86 | 23.6196 | 114.7913 | 557.8855 | 8.1587 | 7.814 | 8.5034 |  | 7.4222 | 7.4061 | 7.4382 |  | 7.422 | 7.4107 | 7.4332 |  | 7.4374 | 7.4271 | 7.4478 |  | 0.020572 | 0.0012595 | -3.3781e-05 | 3.9215e-05 | -4.2209e-06 | -1.8926e-05 | 1.6253e-05 | 2.3326e-06 | -4.8544e-07 | 1.6597e-05 | -7.0839e-05 | 4.3844e-05 | -7.7498e-06 | 4.1538e-07
1 | 5.04 | 25.4016 | 128.0241 | 645.2413 | 8.5008 | 8.1525 | 8.8492 |  | 7.7804 | 7.7645 | 7.7963 |  | 7.7801 | 7.7687 | 7.7916 |  | 7.7944 | 7.7842 | 7.8047 |  | 0.017738 | 0.0018893 | -3.5358e-05 | 3.8993e-05 | -4.1281e-06 | -1.3347e-05 | 4.97e-06 | 5.5825e-06 | -7.1931e-07 | 1.7671e-05 | -7.1125e-05 | 4.245e-05 | -7.315e-06 | 3.8392e-07
1 | 5.22 | 27.2484 | 142.2366 | 742.4753 | 8.8429 | 8.4895 | 9.1963 |  | 8.1452 | 8.1294 | 8.1609 |  | 8.1448 | 8.1331 | 8.1565 |  | 8.1575 | 8.1475 | 8.1676 |  | 0.014904 | 0.002519 | -3.6462e-05 | 3.8438e-05 | -3.9983e-06 | -7.62e-06 | -6.1448e-06 | 8.7261e-06 | -9.4255e-07 | 1.8063e-05 | -6.907e-05 | 3.9801e-05 | -6.6642e-06 | 3.408e-07
1 | 5.4 | 29.16 | 157.464 | 850.3056 | 9.185 | 8.8252 | 9.5447 |  | 8.5164 | 8.5009 | 8.532 |  | 8.516 | 8.504 | 8.5279 |  | 8.5266 | 8.5168 | 8.5365 |  | 0.01207 | 0.0031488 | -3.7093e-05 | 3.7548e-05 | -3.8313e-06 | -1.8533e-06 | -1.6923e-05 | 1.1716e-05 | -1.1516e-06 | 1.7772e-05 | -6.4756e-05 | 3.5965e-05 | -5.812e-06 | 2.8692e-07
1 | 5.58 | 31.1364 | 173.7411 | 969.4754 | 9.5271 | 9.1597 | 9.8945 |  | 8.8941 | 8.8788 | 8.9095 |  | 8.8936 | 8.8813 | 8.9058 |  | 8.9019 | 8.8922 | 8.9117 |  | 0.0092365 | 0.0037786 | -3.725e-05 | 3.6324e-05 | -3.6272e-06 | 3.8445e-06 | -2.7198e-05 | 1.4503e-05 | -1.343e-06 | 1.6814e-05 | -5.8316e-05 | 3.1035e-05 | -4.7774e-06 | 2.2347e-07
1 | 5.76 | 33.1776 | 191.103 | 1100.7531 | 9.8692 | 9.493 | 10.2454 |  | 9.2783 | 9.2632 | 9.2935 |  | 9.2777 | 9.2652 | 9.2902 |  | 9.2834 | 9.2737 | 9.293 |  | 0.0064026 | 0.0044083 | -3.6935e-05 | 3.4767e-05 | -3.3859e-06 | 9.3651e-06 | -3.6802e-05 | 1.7041e-05 | -1.5131e-06 | 1.5224e-05 | -4.9933e-05 | 2.5132e-05 | -3.5845e-06 | 1.5188e-07
1 | 5.94 | 35.2836 | 209.5846 | 1244.9324 | 10.2113 | 9.8251 | 10.5974 |  | 9.669 | 9.6541 | 9.6839 |  | 9.6683 | 9.6555 | 9.6811 |  | 9.6711 | 9.6614 | 9.6807 |  | 0.0035686 | 0.0050381 | -3.6146e-05 | 3.2875e-05 | -3.1076e-06 | 1.46e-05 | -4.5566e-05 | 1.9281e-05 | -1.6584e-06 | 1.3052e-05 | -3.9839e-05 | 1.8403e-05 | -2.262e-06 | 7.3869e-08
1 | 6.12 | 37.4544 | 229.2209 | 1402.8321 | 10.5534 | 10.1563 | 10.9504 |  | 10.0661 | 10.0514 | 10.0808 |  | 10.0654 | 10.0524 | 10.0784 |  | 10.0651 | 10.0553 | 10.0748 |  | 0.00073472 | 0.0056678 | -3.4885e-05 | 3.065e-05 | -2.7922e-06 | 1.9441e-05 | -5.3325e-05 | 2.1175e-05 | -1.7754e-06 | 1.0365e-05 | -2.832e-05 | 1.1023e-05 | -8.4306e-07 | -8.6128e-09
1 | 6.3 | 39.69 | 250.047 | 1575.2961 | 10.8954 | 10.4865 | 11.3044 |  | 10.4697 | 10.4552 | 10.4843 |  | 10.469 | 10.4558 | 10.4821 |  | 10.4655 | 10.4555 | 10.4755 |  | -0.0020992 | 0.0062976 | -3.315e-05 | 2.809e-05 | -2.4396e-06 | 2.378e-05 | -5.9909e-05 | 2.2676e-05 | -1.8604e-06 | 7.2486e-06 | -1.571e-05 | 3.1916e-06 | 6.3412e-07 | -9.3336e-08
1 | 6.48 | 41.9904 | 272.0978 | 1763.1937 | 11.2375 | 10.8158 | 11.6593 |  | 10.8798 | 10.8654 | 10.8942 |  | 10.879 | 10.8658 | 10.8923 |  | 10.8724 | 10.8621 | 10.8827 |  | -0.0049331 | 0.0069274 | -3.0942e-05 | 2.5197e-05 | -2.05e-06 | 2.7508e-05 | -6.5151e-05 | 2.3736e-05 | -1.9101e-06 | 3.8051e-06 | -2.3924e-06 | -4.8623e-06 | 2.1268e-06 | -1.7781e-07
1 | 6.66 | 44.3556 | 295.4083 | 1967.4193 | 11.5796 | 11.1444 | 12.0149 |  | 11.2964 | 11.282 | 11.3108 |  | 11.2956 | 11.2823 | 11.3089 |  | 11.2859 | 11.2753 | 11.2966 |  | -0.007767 | 0.0075571 | -2.8261e-05 | 2.1969e-05 | -1.6232e-06 | 3.0516e-05 | -6.8885e-05 | 2.4308e-05 | -1.9208e-06 | 1.5363e-07 | 1.1197e-05 | -1.2884e-05 | 3.5874e-06 | -2.5929e-07
1 | 6.84 | 46.7856 | 320.0135 | 2188.8924 | 11.9217 | 11.4722 | 12.3712 |  | 11.7194 | 11.7049 | 11.7339 |  | 11.7186 | 11.7054 | 11.7319 |  | 11.7062 | 11.6951 | 11.7172 |  | -0.010601 | 0.0081869 | -2.5106e-05 | 1.8408e-05 | -1.1593e-06 | 3.2697e-05 | -7.0942e-05 | 2.4342e-05 | -1.889e-06 | -3.5693e-06 | 2.4572e-05 | -2.0593e-05 | 4.9639e-06 | -3.3477e-07
1 | 7.02 | 49.2804 | 345.9484 | 2428.5578 | 12.2638 | 11.7994 | 12.7282 |  | 12.1489 | 12.1342 | 12.1636 |  | 12.1482 | 12.1351 | 12.1613 |  | 12.1333 | 12.1218 | 12.1447 |  | -0.013435 | 0.0088166 | -2.1479e-05 | 1.4512e-05 | -6.5836e-07 | 3.3942e-05 | -7.1155e-05 | 2.3792e-05 | -1.8111e-06 | -7.2099e-06 | 3.7198e-05 | -2.768e-05 | 6.1992e-06 | -4.0097e-07
1 | 7.2 | 51.84 | 373.248 | 2687.3856 | 12.6059 | 12.126 | 13.0858 |  | 12.5849 | 12.5699 | 12.6 |  | 12.5842 | 12.5713 | 12.5972 |  | 12.5673 | 12.5556 | 12.579 |  | -0.016269 | 0.0094464 | -1.7378e-05 | 1.0283e-05 | -1.2026e-07 | 3.4143e-05 | -6.9356e-05 | 2.261e-05 | -1.6837e-06 | -1.0597e-05 | 4.8487e-05 | -3.3811e-05 | 7.2317e-06 | -4.5435e-07
1 | 7.38 | 54.4644 | 401.9473 | 2966.3709 | 12.948 | 12.452 | 13.444 |  | 13.0274 | 13.0118 | 13.043 |  | 13.0268 | 13.014 | 13.0395 |  | 13.0085 | 12.9966 | 13.0204 |  | -0.019103 | 0.010076 | -1.2804e-05 | 5.7191e-06 | 4.5494e-07 | 3.3191e-05 | -6.5378e-05 | 2.0747e-05 | -1.5031e-06 | -1.3544e-05 | 5.7803e-05 | -3.8623e-05 | 7.9951e-06 | -4.9112e-07
1 | 7.56 | 57.1536 | 432.0812 | 3266.534 | 13.2901 | 12.7776 | 13.8026 |  | 13.4763 | 13.46 | 13.4927 |  | 13.4758 | 13.4631 | 13.4885 |  | 13.4569 | 13.4449 | 13.4689 |  | -0.021937 | 0.010706 | -7.7573e-06 | 8.2157e-07 | 1.0673e-06 | 3.0977e-05 | -5.9053e-05 | 1.8156e-05 | -1.2658e-06 | -1.5844e-05 | 6.4459e-05 | -4.1729e-05 | 8.4184e-06 | -5.0722e-07
1 | 7.74 | 59.9076 | 463.6848 | 3588.9205 | 13.6322 | 13.1028 | 14.1616 |  | 13.9317 | 13.9145 | 13.949 |  | 13.9314 | 13.9186 | 13.9441 |  | 13.9128 | 13.9008 | 13.9247 |  | -0.024771 | 0.011336 | -2.2372e-06 | -4.4099e-06 | 1.7167e-06 | 2.7395e-05 | -5.0213e-05 | 1.479e-05 | -9.6836e-07 | -1.7276e-05 | 6.7718e-05 | -4.2711e-05 | 8.4256e-06 | -4.9834e-07
1 | 7.92 | 62.7264 | 496.7931 | 3934.6013 | 13.9743 | 13.4275 | 14.521 |  | 14.3936 | 14.3752 | 14.4121 |  | 14.3934 | 14.3804 | 14.4064 |  | 14.3763 | 14.3645 | 14.388 |  | -0.027604 | 0.011965 | 3.756e-06 | -9.9754e-06 | 2.4032e-06 | 2.2334e-05 | -3.8692e-05 | 1.0599e-05 | -6.0712e-07 | -1.7599e-05 | 6.6791e-05 | -4.1129e-05 | 7.9365e-06 | -4.5988e-07
1 | 8.1 | 65.61 | 531.441 | 4304.6721 | 14.3164 | 13.7519 | 14.8808 |  | 14.862 | 14.8423 | 14.8818 |  | 14.8619 | 14.8483 | 14.8756 |  | 14.8475 | 14.8359 | 14.8592 |  | -0.030438 | 0.012595 | 1.0222e-05 | -1.5875e-05 | 3.1269e-06 | 1.5687e-05 | -2.4321e-05 | 5.5376e-06 | -1.7857e-07 | -1.6557e-05 | 6.0839e-05 | -3.6513e-05 | 6.8657e-06 | -3.87e-07
1 | 8.28 | 68.5584 | 567.6636 | 4700.2542 | 14.6584 | 14.0759 | 15.241 |  | 15.3368 | 15.3156 | 15.3581 |  | 15.337 | 15.3223 | 15.3517 |  | 15.3267 | 15.315 | 15.3385 |  | -0.033272 | 0.013225 | 1.7162e-05 | -2.2108e-05 | 3.8876e-06 | 7.3449e-06 | -6.9335e-06 | -4.4346e-07 | 3.2082e-07 | -1.3877e-05 | 4.8974e-05 | -2.8367e-05 | 5.1233e-06 | -2.7461e-07
1 | 8.46 | 71.5716 | 605.4957 | 5122.4939 | 15.0005 | 14.3997 | 15.6014 |  | 15.8182 | 15.7952 | 15.8411 |  | 15.8185 | 15.8022 | 15.8349 |  | 15.8142 | 15.8018 | 15.8266 |  | -0.036106 | 0.013855 | 2.4575e-05 | -2.8676e-05 | 4.6855e-06 | -2.8001e-06 | 1.3639e-05 | -7.3916e-06 | 8.946e-07 | -9.2675e-06 | 3.0257e-05 | -1.6169e-05 | 2.6148e-06 | -1.1733e-07
1 | 8.64 | 74.6496 | 644.9725 | 5572.5628 | 15.3426 | 14.7232 | 15.9621 |  | 16.306 | 16.2811 | 16.3308 |  | 16.3066 | 16.2881 | 16.3251 |  | 16.31 | 16.296 | 16.3239 |  | -0.03894 | 0.014484 | 3.2461e-05 | -3.5577e-05 | 5.5204e-06 | -1.4857e-05 | 3.7564e-05 | -1.5355e-05 | 1.5463e-06 | -2.4198e-06 | 3.6984e-06 | 6.3161e-07 | -7.5922e-07 | 9.0475e-08
1 | 8.82 | 77.7924 | 686.129 | 6051.6575 | 15.6847 | 15.0464 | 16.323 |  | 16.8002 | 16.7734 | 16.8271 |  | 16.8012 | 16.7799 | 16.8224 |  | 16.8144 | 16.7976 | 16.8312 |  | -0.041774 | 0.015114 | 4.082e-05 | -4.2812e-05 | 6.3925e-06 | -2.8933e-05 | 6.5008e-05 | -2.4381e-05 | 2.2795e-06 | 6.9912e-06 | -3.1743e-05 | 2.2611e-05 | -5.1028e-06 | 3.5469e-07
1 | 9 | 81 | 729 | 6561 | 16.0268 | 15.3694 | 16.6842 |  | 17.301 | 17.2719 | 17.33 |  | 17.3022 | 17.2777 | 17.3268 |  | 17.3276 | 17.3067 | 17.3485 |  | -0.044608 | 0.015744 | 4.9652e-05 | -5.0382e-05 | 7.3017e-06 | -4.5138e-05 | 9.6139e-05 | -3.4517e-05 | 3.0977e-06 | 1.9308e-05 | -7.7156e-05 | 5.0371e-05 | -1.0525e-05 | 6.8146e-07